Particle Measurement Programme (PMP) Heavy-duty Inter-laboratory Correlation Exercise (ILCE_HD) Final Report

Jon Andersson, Athanasios Mamakos, Barouch Giechaskiel, Massimo Carriero, Giorgio Martini

January 2010

Contents

1.	INTRODUCTION	10

2.	OBJECTIVES	11

3.	NATURE AND SCOPE OF THE ILCE_HD	11

3.1	Background to the PMP	11

3.2	PMP Phase 3 Inter-laboratory Exercise for Light-duty Vehicles	11

3.3	Brief Overview of the Inter-laboratory Correlation Exercises for Heavy-duty Engines	13

3.4	Test Engine and Emissions Control System	14

3.4.1	CRT: Pt-based Oxidation Catalyst and Wall-flow DPF	15

3.5	Fuel and Lubricant	16

3.6	Gaseous Emissions Measurement Systems	16

3.7	Dilution Approaches	17

3.7.1	Principles of the Dilution Systems	17

3.8	Particulate Mass Measurements	20

3.8.1	Full Flow Measurements	20

3.8.2	Partial flow Measurements	21

3.9	Particle Number Measurement Systems	22

3.9.1	Principles of the Measurement System	22

3.9.2	Golden Particle Measurement Systems	22

3.9.3	Other PMP type systems	24

3.9.4	PMP like systems:	24

3.9.5	Additional Instrumentation	25

3.10	Test Programme	25

3.10.1	Participating Laboratories	25

3.10.2	Summary of Preliminary Experiments	26

3.10.3	Daily Protocol	34

4.	STATISTICAL ANALYSES	35

4.1	Definitions	35

4.1.1	Basic statistical concepts	35

4.1.2	Intra- And Inter-Laboratory Variance Estimators	36

4.1.3	Statistical outliers	36

5.	CALIBRATION AND VALIDATION	38

5.1	Mass systems	38

5.2	Number systems	38

5.2.1	SPCS Calibration	39

5.2.2	3010D Calibration	44

5.2.3	Validation Exercises	44

6.	EMISSIONS RESULTS	45

6.1	Full Flow and Partial Flow PM	45

6.1.1	PM - Repeatability	45

6.1.2	PM - Reproducibility	47

6.1.3	PM – Filter Weights, Background Levels and Background Subtraction	48

6.1.4	PM – Emissions Levels	51

6.1.5	PM – Daily Trends	56

6.1.6	Overview of PM Results	57

6.2	Full Flow and Partial Flow PN	57

6.2.1	PN – Tunnel background Levels	57

6.2.2	PN - Repeatability	59

6.2.3	PN - Reproducibility	67

6.2.4	PN – Emissions Levels	68

6.3	Gaseous Emissions	75

6.3.1	General Observations	75

6.3.2	Inter-Laboratory Variability	75

6.4	Long-Term Trends in Engine Operation	Error! Bookmark not defined.

6.4.1	Regulated Gases	Error! Bookmark not defined.

6.4.2	Particle Number Emissions	Error! Bookmark not defined.

7.	ADDITIONAL TESTWORK OVERVIEW	Error! Bookmark not defined.

7.1	Mass Contributed by EC and by Particles	Error! Bookmark not defined.

7.2	Further Filter Media Evaluations	Error! Bookmark not defined.

7.3	Particle Number Measurements from PFDS at Constant Dilution Factors	Error! Bookmark not defined.

8.	COMPARISONS OF MEASUREMENT SYSTEMS	Error! Bookmark not defined.

8.1	Comparison of Particle Number Measurements from CVS and PFDS	Error! Bookmark not defined.

8.2	Comparison of Particulate Mass Systems	Error! Bookmark not defined.

8.3	Mass vs. Number full flow	Error! Bookmark not defined.

8.4	Mass Vs. Number Partial Flow	Error! Bookmark not defined.

8.5	Comparisons of PMP type and PMP like systems	Error! Bookmark not defined.

8.5.1	PMP type and PMP- Like Systems	Error! Bookmark not defined.

8.5.2	Particle Losses	Error! Bookmark not defined.

8.6	Results from Golden Engine equipped with DPF	Error! Bookmark not defined.

8.6.1	PMP type systems	Error! Bookmark not defined.

8.6.2	PMP like systems	Error! Bookmark not defined.

8.7	Engine out and open flow emission levels	Error! Bookmark not defined.

8.8	Gaseous Emissions	Error! Bookmark not defined.

9.	DISCUSSION AND OVERVIEW	Error! Bookmark not defined.

9.1	Particulate Measurements	Error! Bookmark not defined.

9.1.1	Validation Exercises	Error! Bookmark not defined.

9.1.2	Tunnel Background Mass Levels - CVS	Error! Bookmark not defined.

9.1.3	Tunnel Background Mass Levels - PFDS	Error! Bookmark not defined.

9.1.4	Emissions Levels	Error! Bookmark not defined.

9.1.5	Tunnel Background Subtraction	Error! Bookmark not defined.

9.1.6	Partial Flow or Full Dilution	Error! Bookmark not defined.

9.2	Particle Number Measurements	Error! Bookmark not defined.

9.2.1	Golden Systems – PCRF Correction	Error! Bookmark not defined.

9.2.2	Particle Number Tunnel backgrounds	Error! Bookmark not defined.

9.2.3	Real-time Particle Emissions Elevated at Cold Start	Error! Bookmark not defined.

9.2.4	PN Emissions Levels Compared with Backgrounds	Error! Bookmark not defined.

9.2.5	PN Emissions levels	Error! Bookmark not defined.

9.3	Simultaneous Particle Number and Particulate Mass Measurements	Error! Bookmark not defined.

9.3.1	Required Corrections	Error! Bookmark not defined.

9.4	Repeatability and Reproducibility	Error! Bookmark not defined.

9.4.1	PM and PN – Repeatability	Error! Bookmark not defined.

9.4.2	PM and PN – Reproducibility	Error! Bookmark not defined.

9.4.3	Repeatability and Reproducibility Overview	Error! Bookmark not defined.

9.5	Elemental carbon and non-volatile particle contributions to PM	Error! Bookmark not defined.

9.6	Filter Media Effects	Error! Bookmark not defined.

9.7	Other PMP type / like Systems	Error! Bookmark not defined.

9.8	Should Solid Particles <23nm be Considered for European PN Legislation?	Error! Bookmark not defined.

9.9	Achievable PN Limit for this Engine	Error! Bookmark not defined.

9.10	An Alternative Approach to Sampling for PN Measurements	Error! Bookmark not defined.

10.	CONCLUSIONS	Error! Bookmark not defined.

10.1	Engine Operation	Error! Bookmark not defined.

10.2	Measurement Systems	Error! Bookmark not defined.

10.3	PM Emissions	Error! Bookmark not defined.

10.3.1	Levels	Error! Bookmark not defined.

10.3.2	Repeatability / Reproducibility	Error! Bookmark not defined.

10.4	PN Emissions	Error! Bookmark not defined.

10.4.1	Levels	Error! Bookmark not defined.

10.4.2	Repeatability / Reproducibility	Error! Bookmark not defined.

10.5	Relationships Between Measurement Approaches	Error! Bookmark not defined.

10.5.1	Full Vs Partial Flow Mass	Error! Bookmark not defined.

10.5.2	Full Vs Partial Flow Number	Error! Bookmark not defined.

10.5.3	Mass Vs Number	Error! Bookmark not defined.

10.6	Other PMP type / Like Systems	Error! Bookmark not defined.

10.7	General Conclusions	Error! Bookmark not defined.

11.	RECOMMENDATIONS	Error! Bookmark not defined.

List of Figures

Figure 1: Typical Engine and Emissions Control System Installations	14

Figure 2: Schematic of Exhaust and Emissions Control System Layout	15

Figure 3: Schematic of Full Flow, Double Dilution System for PM Measurements	17

Figure 4: Schematic of Partial Flow Dilution System for PM Measurements	18

Figure 5: PFDS – Principle of Dilution	19

Figure 6: Flow schematic of the prototype SPCS	23

Figure 7: PM Sampling Set-up for Experiment 1	28

Figure 8: Experiment 1 - Background and Sampling Parameters for PM	28

Figure 9: Minimal Impact of Heated Line on PN Results	30

Figure 10: Minimal Impact of Cyclone on PN	30

Figure 11: DPF Preconditioning –Regeneration (Mode 10) and Fill (Mode 7)	32

Figure 12: Stable Mode 7 Selected for the Continuity Protocol	32

Figure 13: Particle Number Emissions <23nm	33

Figure 14: Weighing Room Validation Exercises - Ricardo	38

Figure 15: Apparatus for the Evaluation of Volatile Particle RE	39

Figure 16: VPR RE of 99.99% for 30nm n-C40 particles	40

Figure 17: Apparatus for Evaluation of Solid Particle Penetration (Polydisperse)	41

Figure 18: Size Distributions (downstream dilution correction of 1050 applied)	41

Figure 19: Penetration at Different Dilution Ratios (Polydisperse)	41

Figure 20: Apparatus for Evaluation of Solid Particle Penetration (Monodisperse)	42

Figure 21: SPCS Penetration of 30nm Particles	42

Figure 22: SPCS Penetration of 50nm Particles	43

Figure 23: SPCS Penetration of 100nm Particles	43

Figure 24: Dilution Ratio Calibration – Fixed Secondary Dilution	44

Figure 25: Dilution Ratio Calibration – Fixed Primary Dilution	44

Figure 26: Repeatability of the CVS PM Method	46

Figure 27: Repeatability of the PFDS PM Method	47

Figure 28: Reproducibility of PM Methods	48

Figure 29: Sampled Filter Masses – Various Cycles, Both Measurement Systems	48

Figure 30: Filter Masses - Samples and Tunnel Backgrounds Compared	49

Figure 31: Tunnel Background and Sample PM levels - CVS	50

Figure 32: Background and Sample PM levels - PFDS	51

Figure 33: Maximum and Minimum Ranges of Emissions – CVS PM	54

Figure 34: Maximum and Minimum Ranges of Emissions – PFDS PM	55

Figure 35: CVS PM Emissions through the Test Sequence – All labs	56

Figure 36: PFDS PM Emissions through the Test Sequence – All labs	57

Figure 37: Tunnel background Impacts PN Results – CVS Systems at 3 Labs	58

Figure 38: Tunnel background Effect on PN Results Minimal– PFDS Systems at 3 Labs	59

Figure 39: Particle Number Tunnel backgrounds – 3 Labs; CVS and PFDS Facilities	59

Figure 40: Real-time Data – Cycle-to-cycle Repeatability, Cold WHTC	61

Figure 41: Real-time Data – Cycle-to-cycle Repeatability, Hot WHTC	62

Figure 42: Real-time Data – Cycle-to-cycle Repeatability, Hot WHSC	63

Figure 43: Passive Regeneration During the WHSC	64

Figure 44: Real-time Data – Cycle-to-cycle Repeatability, Hot ESC	65

Figure 45: Repeatability of Particle Number Measurement Systems	67

Figure 46: Reproducibility of Particle Number Measurement Systems	68

Figure 47: Maximum and Minimum Ranges for CVS-measured PN	70

Figure 48: CVS PN Results – Effect of Eliminating Outliers	71

Figure 49: Maximum and Minimum Ranges for PFDS-measured PN	73

Figure 50: All Labs’ Mean of Means PN Results - PFDS Systems	74

Figure 51: CO2 – Consistency Across The Validation Exercise	Error! Bookmark not defined.

Figure 52: NOx – Consistency Across The Validation Exercise	Error! Bookmark not defined.

Figure 53: PN from cold-WHTC: Consistency Through The Test Programme	Error! Bookmark not defined.

Figure 54: Real time PN Production from the CVS- Cold WHTC (JRC#1)	Error! Bookmark not defined.

Figure 55: Real time PN Production from the CVS- Cold WHTC (JRC#2)	Error! Bookmark not defined.

Figure 56: Fraction of PM Mass Contributed by Solid Particles (>300°C)	Error! Bookmark not defined.

Figure 57: Fraction of Mass Contributed by EC	Error! Bookmark not defined.

Figure 58: PM results obtained using TX40 and TEFLO filters – JRC CVS	Error! Bookmark not defined.

Figure 59: PM results obtained using TX40 and TEFLO filters – JRC PFDS	Error! Bookmark not defined.

Figure 60: Cycle Averaged PN Results: Proportional Vs. Constant Dilution	Error! Bookmark not defined.

Figure 61: PN from Proportional and Constant Dilution Sampling Compared	Error! Bookmark not defined.

Figure 62: Correlation of PN from Proportional and Constant Dilution Sampling	Error! Bookmark not defined.

Figure 63: PM Sampled From Proportional Dilution and Constant Dilution	Error! Bookmark not defined.

Figure 64: Correlations between PN Measured from CVS and PFDS Systems	Error! Bookmark not defined.

Figure 65: Correlations between Real-time PN from CVS and PFDS Systems	Error! Bookmark not defined.

Figure 66: Percentage Differences in PFDS Emissions Compared to CVS – All Labs	Error! Bookmark not defined.

Figure 67: Correlations between PM Measured from CVS and PFDS Systems	Error! Bookmark not defined.

Figure 68: PFDS PM results: % Difference to CVS PM Data	Error! Bookmark not defined.

Figure 69: Relationship between Mass and Number Measurements (CVS)	Error! Bookmark not defined.

Figure 70: Relationship between Mass and Number Measurements (PFDS)	Error! Bookmark not defined.

Figure 71: No Correlation Exists PM & PN Even in a Very Low Emissions PFDS	Error! Bookmark not defined.

Figure 72: Correlation between the GPMS and other PMP type systems	Error! Bookmark not defined.

Figure 73: Comparison of real time PN in the CVS from the GPMS and from a dual ejector plus ET system over Cold and Hot WHTCs	Error! Bookmark not defined.

Figure 74: Correlation between the GPMS and PMP like systems	Error! Bookmark not defined.

Figure 75: Correlation between the GPMS and other PMP type systems at engine out emission levels	Error! Bookmark not defined.

Figure 76: Correlation between the GPMS and PMP like systems at engine out emission levels	Error! Bookmark not defined.

List of Tables

Table 1: Test Engine and Specification	14

Table 2: Exhaust System Installation Variability at the Test Laboratories	15

Table 3: Fuel specifications	16

Table 4: Lubricant Properties	16

Table 5: Principal Differences Between CVS Systems – Test Labs	21

Table 6: Principal Differences Between PFDS Systems – Test Labs	22

Table 7: Test Laboratories and Timeline	26

Table 8: Matrix for Emissions Testing	34

Table 9: Limits of Detection – PM Methods at JRC and EMPA (mg/kWh)	51

Table 10: Maximum and Minimum Ranges for CVS-measured PN	69

Table 11: Maximum and Minimum Ranges for PFDS-measured PN	72

Table 12: Contribution of soot and solid airborne particles on the PM collected on filters as determined by the Soot Sensor and the DMM, respectively	Error! Bookmark not defined.

Table 13: Particle penetrations through the various PMP type / like systems used in the PMP HD validation exercise.	Error! Bookmark not defined.

Table 14: Percentage difference between the different CPCs employed and the Golden CPC of SPCS19.	Error! Bookmark not defined.

Table 15: Percentage differences in the number concentrations measured with the GPMS, PMP type and PMP-like systems	Error! Bookmark not defined.

List of Appendices

Appendix 1: ILG_HD	Error! Bookmark not defined.

Appendix 2: Certification of PNC Performance	Error! Bookmark not defined.

Executive Summary

The Heavy Duty Inter-Laboratory Correlation Exercise has conducted testing at 5 test laboratories in the Europe in order to demonstrate the practicality, robustness, repeatability and reproducibility of the particle emissions measurement techniques proposed by the Particle Measurement Programme (PMP). The exercise involved testing a Golden Engine (a Euro III Iveco Cursor 8, equipped with a wall-flow Diesel Particulate Filter), at all participating laboratories to allow the inter-laboratory reproducibility of measurements to be assessed. Each laboratory tested the engine over multiple repeats of the heavy duty World Harmonised Transient Cycle (both cold and hot start) and World Harmonised Steady state Cycle (WHSC) as well as current EU regulatory cycles the European Transient Cycle (ETC) and European Steady state Cycle (ESC). Measurements of solid particle number emissions, particulate mass and regulated gaseous emissions were taken over each test. Particle emissions measurements were taken from both full flow (CVS) and partial flow (PFDS) dilution systems at each laboratory. Two ‘Golden’ particle number measurement systems were circulated between the test laboratories, one for use in CVS measurements one for use in PFDS measurements. In addition laboratories made particle number measurements using several alternative, PMP type systems to compare the performance of different measurement systems. The Golden Measurement systems performed reliably at all laboratories and agreed within 5% when making measurements in parallel.

Particulate mass (PM) emissions levels from the Golden Engine using CVS sampling systems were below 6mg/kWh across all test cycles after exclusion of outlying test results. High tunnel background contributions in some laboratories CVS systems resulted in higher outlying results. PFDS systems returned slightly lower PM results (below 4mg/kWh). In this exercise tunnel background PM measurements were generally found to be similar to engine measurements, although ESC cycle results sampled from PFDS systems could be discriminated from tunnel background levels by all laboratories. PM measurement repeatability from PFDS was 20-30% for all test cycles and rather higher 35-56% for CVS measurements after exclusion of outliers. Reproducibility between laboratories was 35-45% for PFDS measurements and 35-55% for CVS measurements.

Particle number (PN) emissions levels from the Golden Engine varied significantly from cycle to cycle. The cold start WHTC gave the highest PN levels, approximately 4x1011/kWh from both CVS and PFDS dilution systems. At these levels tunnel background PN concentrations did not significantly influence measurements. Hot start WHTC and ETC cycles gave PN levels around 5-9x109 /kWh, steady state cycles gave higher results (2-3x1010 /kWh on the WHSC and 6-8x1010 /kWh on the ESC) possibly due to higher exhaust temperatures resulting in some passive regeneration and resulting reduction in filtration efficiency as the soot cake on the Diesel Particulate Filter is reduced. On these test cycles tunnel background levels were found to have a significant impact in the case of some laboratories CVS systems, PFDS tunnel background levels however were significantly lower in did not influence PN results. Where tunnel background concentrations were low correlation between CVS and PFDS measurements was excellent.

PN repeatability levels across the different test cycles ranged from 20-60% for CVS sampling with best repeatability being on the cold WHTC (where PN levels were highest) and worst on the WHSC where partial passive regeneration of the Diesel Particulate Filter may result in less stable PN emissions from the engine. PFDS repeatability ranged from 20-70%, with best and worst results again on cold WHTC and WHSC respectively, however, as noted above, on the lower emissions cycles PFDS measurements were less influenced by tunnel background levels than was the case for CVS measurements. PN reproducibility between laboratories was generally similar to repeatability. Across the different test cycles PN reproducibility ranged from 30-80% for CVS sampling and 50-86% for PFDS sampling.

Results from alternative PN measurement systems conforming to PMP principles generally correlated to the Golden System measurements within 15% although some systems showed greater offsets. Additional experiments conducted during the exercise for investigative purposes showed that, contrary to results from some US research, reported concentrations of solid particles of less than 23nm diameter were low relative to those of larger than 23nm particles. This confirms the suitability of the 23nm lower size cut-off for the particle number counter.

The results of this exercise demonstrate that the PM measurement method is suitable to confirm that engine emissions are below 10mg/kWh. PM measurements of an engine equipped with an efficient wall-flow Diesel Particulate Filter were similar across all test cycles. In this exercise PM measurements could not generally be discriminated from tunnel background PM measurements. The PN emissions measurement method was able to discriminate between the emissions levels on different test cycles of an engine equipped with an efficient wall-flow Diesel Particulate Filter. PN was also able to discriminate engine emissions from tunnel background levels in this exercise except in the case of high tunnel background sampling systems during testing on cycles with lower emissions levels.

Particle Measurement Programme (PMP) Heavy-duty Inter-laboratory Correlation Exercise (ILCE_HD) Final Report

# INTRODUCTION

The effect of exhaust emissions from road vehicles on public health has long been a concern. Legislation limiting the pollutant emissions of new vehicles is well established in many regions of the world. One emission of special concern is particulate matter. In vehicle exhaust this consists of tiny solid particles and liquid droplets ranging in size from a few nanometres to up to around one micrometre in diameter. Current legislative emissions standards regulate particle emissions in terms of the total mass of particulate matter emitted per kilowatt hour. This is effective at controlling emissions of larger size particles, but particles at the smaller end of the size range contribute little to the total mass of particulate matter emitted.

There is a growing consensus amongst health experts that particles in the ultrafine (<100nm diameter) size range may be those which are having the greatest adverse effect on human health. The main driver behind Particle Measurement Programme (PMP) is the impact of particles on human health. The PMP has no medical expertise and does not seek to pre-judge the advice that may emerge from medical experts with respect to the most crucial particle characteristics affecting human health. Nonetheless, current medical opinion suggests that reductions in particle emissions will lead to improved air quality and health and the PMP has therefore moved forward on the basis of the precautionary principle. This and the potential limitations of current regulatory procedures at forcing technology that would control these particle emissions led to the setting up of the PMP as a Working Group of the UN-ECE GRPE. PMP is essentially a collaborative programme of Government sponsored research projects. However the Working Group, chaired by the UK, exists to co-ordinate the research and ensures that the programme is open to contributions from a wider audience. National Governments, individual laboratories, exhaust aftertreatment, automotive industry and fuel industry representatives have all provided significant input to the programme.

The mandate given to the PMP Working Group by GRPE was to develop new particle measurement techniques to complement or replace the existing particulate mass measurement, with special consideration to measuring particle emissions at very low levels. These techniques should include a detailed specification of test procedures and equipment, be suitable for Light Duty Vehicle and Heavy Duty Engine type approval testing and be suitable for use in transient testing. Since, within the EU, type approval testing to demonstrate compliance with emissions standards involves a limited number of tests which could take place at one of many laboratories, good repeatability and reproducibility from laboratory-to-laboratory are key requirements for regulatory measurement techniques. PMP has therefore sought to demonstrate the repeatability and reproducibility of the proposed techniques. PMP was also tasked with accumulating data on the performance of a range of engine/vehicle technologies when tested according to the proposed procedures.

# OBJECTIVES

# NATURE AND SCOPE OF THE ILCE_HD

## Background to the PMP

In 2001, the French, German, Netherlands, Swedish and United Kingdom Governments agreed to a collaborative programme aimed at developing new methods and procedures to facilitate the control of ultrafine particles within a regulatory framework. This programme was designed to deliver a regulatory procedure that would either replace or complement the existing procedure used for particulate mass measurement.

The resulting Particle Measurement Programme (PMP) working group, chaired by the United Kingdom’s Department for Transport, operates under the auspices of the UN-ECE, where the government of Switzerland joined the consortium. Japanese and Korean governments have also contributed.

The PMP working group devised a three-phased approach to the PMP Programme.

In the first two phases of the programme, a wide range of measurement instruments and sampling systems were assessed over standard regulatory tests:

In the PMP Phase 1 study, measurement systems addressing several key particle properties including mass, number, active surface and chemistry were evaluated along with appropriate dilution methods, sample conditioning and consideration of cost and logistical aspects.

Phase 2 subjected the best performing systems from Phase 1 to more rigorous evaluations.  Aims were to confirm the results of Phase 1 and determine fundamental levels of repeatability within a single laboratory during a variety of steady state and transient tests with both engine-out and post-DPF exhaust aerosols. The testing from Phase 2 enabled the conclusions that a revised filter mass measurement method and a particle number method both, based upon sampling from a standard dilution system, best met the original objectives of the programme. The two recommended systems were:

A filter method based broadly upon those currently used in Europe and the US and that proposed for the US for 2007 type approvals

A particle number method using a Particle Counter, a selected size range and sample pre-conditioning to eliminate volatile particles

Draft revised versions of the light-duty vehicle (DR83 []) and heavy-duty engine (DR49 []) particulate regulatory sampling annexes were prepared from the existing regulatory documents: R83 [] and R49 [].

The new documents integrated the PMP particulate and particle number approaches into the existing regulatory framework and also formed the bases for two test protocol documents written as laboratory guides for testing. The PMP Phase 3 “Inter-Laboratory Correlation Exercises” (ILCE) for light-duty vehicles (ILCE_LD) and heavy-duty engines (ILCE_HD) then commenced with the light duty vehicles’ exercise.

## PMP Phase 3 Inter-laboratory Exercise for Light-duty Vehicles

The light duty vehicles’ exercise circulated a Euro 4 light-duty Diesel vehicle equipped with an OEM fit Diesel Particulate Filter (DPF) plus a reference “Golden” particle measurement system (GPMS) between laboratories. In addition, each lab was invited to employ other particle measurement systems constructed to meet the design criteria of the GPMS, and to test other Euro 4 vehicles. Testing followed the procedures described in the inter-laboratory guide for light-duty vehicles (ILG_LD) [] and comprised the measurement of regulated gaseous emissions, particulate mass and particle number from repeat NEDC tests. To ensure maximum consistency of testing between laboratories, the Golden Engineer and project manager visited the participating labs to advise on facility modifications, how to undertake the test protocols and installation and operation of the GPMS. Low sulphur fuel and lubricant from the same batches were also used at all labs.

The ILCE_LD has now completed, with the final report published in June 2007 [] and extended data analyses published in the scientific literature [], [].

The general conclusions of the ILCE_LD are presented below:

- The revised PMP mass method provides repeatable measurements at well below 2.5 mg/km, but the method collects a large gaseous volatile fraction that may be 20 times the mass of the solid particles collected.

- Both mass and number measurement approaches appear to have detection limits low enough to discriminate between a highly efficient wall-flow DPF equipped Diesel and non-DPF equipped Diesel vehicles. In this testing, the mass method proved unable to discriminate a porous wall-flow DPF from a more efficient one.

- The PMP Particle Number method proved to be less variable than the PMP mass method for Euro-4 non-DPF diesel cars, with repeatability levels from 6 vehicles at ≥5%.

- Comparing the lowest emissions of the non-DPF Diesels and the highest emissions of the efficient wall-flow DPF equipped Diesels, the number method showed a difference of >300 times and the mass method a difference of ~18 times. This can be expressed as a difference in discriminating power approximately 20 times greater for the number method than for the mass method.

- Mass and number measurement equipment presented no significant functional challenges during the 2 year programme. Minor maintenance issues did occur but these were dealt with as normal service issues.

- The PMP number method presents improvements over the PMP mass method in terms of limit of detection, accuracy, discrimination power and variability when measuring a stable particle source. For these reasons, the number method is a superior alternative to the existing or a revised mass method for future regulatory procedures.

The conclusions of the final report, and consultations with stakeholder groups including national Governments, the European Commission, the automotive industry, Tier 1 suppliers and the test houses were used to finalise a new Annex for the R83 which introduced the particle number procedure for certification testing. Modifications to the particulate mass measurement procedure were also integrated. The new procedures came into force with the official publication of the procedures during February 2009.

## Brief Overview of the Inter-laboratory Correlation Exercises for Heavy-duty Engines

Following the successful completion of the ILCE_LD, the PMP working group determined the scope of the heavy-duty exercise. This essentially comprises 3 parts:

- Investigative work to develop and finalise a robust inter-laboratory guide for heavy-duty engines testing. Experiments included identifying background PM and PN levels, effects of different filter media, impacts of filter face velocity changes, exhaust and engine preconditioning effects, comparisons of different particle number systems and investigation of the golden instruments. .A full report of the experimental work has been published previously [,] and an overview is given in Section 3.10.2. The final inter-laboratory guide [] is included in this report as Error! Reference source not found..

- The validation exercise: analogous to the ILCE_LD, this programme investigated particle number repeatability and reproducibility by transporting a Golden engine to each test laboratory in turn. Along with the engine, two Golden Particle Measurement Systems (GPMS) were shipped to permit particle number measurements to be made simultaneously from both full flow constant volume sampler (CVS) dilution and partial flow dilution (PFDS) systems. As in the ILCE_LD, participating laboratories were also invited to test their own particle measurement systems, or other commercially available particle numbers systems. PM and gaseous emissions were also measured. The Golden Engineer and the project manager ensured that participating labs correctly followed the measurement protocols defined in the inter-laboratory guide. Low sulphur fuel and lubricant from the same batches were used at all labs. The participating labs were JRC (Ipsra, Italy), AVL_MTC (Sweden), Ricardo (UK), UTAC (France), and EMPA (Switzerland). Each test laboratory was funded by its respective national government. JRC undertook duplicate measurement campaigns – at the start and end of testing - to monitor consistency of emissions through the programme. Testing in the validation exercise completed at JRC in October 2009.

- The round robin exercise: adhering to the principles of an automotive industry round-robin exercise, this programme is complementary to the validation exercise, but subtly different. Its objective is purely the evaluation of particle number repeatability and reproducibility using different measurement systems. In the round robin, a reference engine is circulated, but each lab uses its own particle number systems from full or partial flow dilution tunnels. All labs will use fuel and lubricant of the same types (but not necessarily from the same batches). PM and regulated gaseous emissions will also be measured. Laboratories from the EU, Japan, Korea and Canada are all participating in the programme. While testing in the round-robin exercise has completed at several laboratories, work is on-going, with completion anticipated during 2011.

This report describes the procedures, results and conclusions of the PMP Phase 3 validation exercise in detail. Once it has completed, the round-robin exercise will be reported separately.

## Test Engine and Emissions Control System

The engine employed in the test programme was a series production IVECO Cursor 8 engine in EU III specification; a modern 7.8 litre, 6-cylinder engine used in heavy-duty vehicle and bus applications. Further details can be found in Table 1.

Table 1: Test Engine and Specification

| IVECO Cursor 8 (Euro 3) |
| --- |
| Details | 7.8 l, 6 cylinder, 4 valves/cylinder |
| Compression ratio | 17:1 |
| Maximum power | 295 kW @ 1900 to 2400 rpm |
| Maximum torque | 1280 Nm @ 1000 to 1900 rpm |
| After-treatment | Continuous Regenerating Trap (CRT) |
| Oxicat | Pt-based: 10.5x3" catalyst section; approx 4.25 litres. |
| DPF | Wall-flow DPF: 11.25x14"; approx 24 litres |

Figure 1 (below left) illustrates a typical installation of the engine, in this case the first test laboratory, JRC. The right-hand part of Figure 1 illustrates the exhaust system layout at JRC which was used to devise a set of benchmark dimensions for installations at all the other test laboratories.

A guide to installation and commissioning was supplied with the test engine and a support engineer visited each test laboratory to facilitate these processes. After testing at JRC this guide was updated to include the exhaust system layout and sampling positions for the PFDS, raw gas analysers and temperature and pressure sensors. A schematic representation of the engine and exhaust layout is given in Figure 2.

Figure 1: Typical Engine and Emissions Control System Installations

A degree of variability, due to the constraints of test cell size and orientation, was expected during engine, exhaust system and PM / PN measurement system installations. Differences are summarised in Table 2. Generally, differences between labs were small and are not expected to have impacted results.

Figure 2: Schematic of Exhaust and Emissions Control System Layout

Numbers in Figure 2 indicate component lengths (and diameters in parentheses) in cm.

Table 2: Exhaust System Installation Variability at the Test Laboratories

| Sampling Dimensions :length in cm diameter (cm) | AVL MTC | JRC | Ricardo | UTAC | EMPA |
| --- | --- | --- | --- | --- | --- |
| Engine-CRT | 250 (15) | 270 (15) | 165 (15) | 100 (15) | 299 (10) |
| CRT-PFDS | 700 (15) | 500 (15) | 395 (15) | 350 (15) | 934 (12.5) |
| CRT-CVS | 1100 (15) | 950 (15) | 930 (15) | 750 (15) | 1469 (12.5) |
| CRT-CVS Insulated | 1100 (15) | 600 (15) | 200 (15) | 450 (15) | 1045 (12.5) |
| PFDS-SPCS20 | 150 | 150 | 400 | 150 | 320 |
| CVS sampling point – CVS mixing point | 500 (50) | 470 (47) | 500 (45) | 575 (45) | 470 |
| CVS-SPCS19 | 400 | 400 | 360* | 400 | 202 |

* Ricardo used a heated line at 47°C to extend the 1m Horiba sampling line to ~4m.

### CRT: Pt-based Oxidation Catalyst and Wall-flow DPF

The DPF employed in the test programme was a cordierite wallflow filter of approximately 24 litres volume and originally supplied by Johnson-Matthey. The ratio of DPF volume to engine size is therefore approximately 3, which is larger than the 1.5 to 2.5 typically employed in current HD applications.

The DPF is preceded in the exhaust system by a close-canned Pt-based oxidation catalyst (Eminox) of approximately 4.25 litres volume.

## Fuel and Lubricant

Fuel and lubricant were supplied to the PMP programme by members of Concawe.

The test fuel was provided by Total, who isolated a large batch of the certification reference fuel RF06-03 and nominated it RF06-03-PMP. Participating labs purchased quantities of this batch directly from the supplier. This fuel fully complies with Annexes 3 and 4 of Directive 2003/17/EC describing fuel specifications to be employed after 1st January 2009 (i.e. sulphur content of < 10 ppm). Selected properties are given in Table 2 and the detailed specifications can be found on the final page of Error! Reference source not found..

Table 3: Fuel specifications

| Properties | Units | Value |
| --- | --- | --- |
| Cetane Number | [-] | 53.1 |
| Density | [kg/m3] | 834.9 |
| Sulphur | [ppm] or [mg/kg] | 7 |
| Polycyclic aromatics | [%] by mass | 5.1 |

The test lubricant (Table 4: Lubricant Properties) was a BP Vanellus E8 fully synthetic, 5W/30 PAO (polyalphaolefin) based oil with <0.2% sulphur content. Defined oil change and conditioning procedures were employed at each laboratory to standardise oil conditioning and eliminate this as a source of variability in the results.

Table 4: Lubricant Properties

| Density @ 15°C | 0.860kg/litre |
| --- | --- |
| Kinematic viscosity @100°C | 12.03mm2/s |
| Viscosity Index | 163 |
| Viscosity CCS @ -30°C | 5260 CP |
| Total Base Number | 15.9 mg KOH/g |
| Sulphated Ash | 0.19% |

## Gaseous Emissions Measurement Systems

During emissions tests at all laboratories selected gaseous emissions were measured on a continuous basis from both raw and diluted exhaust. In addition, some laboratories supplied cumulative 'bagged' sample results.

Raw exhaust samples were drawn directly from the exhaust line, while diluted samples and bagged analyses were made from the full-flow dilution system.

Regulated exhaust gases and their methods of analysis are given below:

- Total hydrocarbons (THC): performed using a heated Flame Ionisation Detector (FID)

- Oxides of Nitrogen (NOx): conducted using a Chemiluminescence Analyser (CLA). CLA detects photons that are emitted by excited NO2 molecules generated in the instrument reaction chamber from NO. Excited NO2 emits photons of a specific wavelength.  The light generated in the reaction is proportional to the NO present in the sample. All the NO2 in the sample gas is reduced to NO prior to the reaction chamber. The combined concentration of NO+NO2 is measured.  As most oxides of nitrogen are generally in one of these two forms, this measurement is expressed as NOx

- Carbon Monoxide (CO) using a Non-Dispersive Infra-red (NDIR) instrument

- Though currently unregulated, carbon dioxide emissions were also measured: using a Non-Dispersive Infra-red (NDIR) instrument.

During the validation exercise, instrumentation provided by the following analyser suppliers was used for gaseous emissions analysis:

- Horiba

- AVL (both own branded and Pierburg)

## Dilution Approaches

### Principles of the Dilution Systems

In Europe since the implementation of Euro IV legislation, two dilution approaches have been considered equivalent for the measurement of particulate mass during certification testing. These full and partial flow dilution approaches are shown in schematic form in Figure 3 and Figure 4 respectively.

Figure 3: Schematic of Full Flow, Double Dilution System for PM Measurements

Figure 4: Schematic of Partial Flow Dilution System for PM Measurements

#### Full flow dilution

In a full flow dilution system, the entire exhaust is sampled and diluted but the total flow through the dilution system is maintained at a constant level. This is known as constant volume sampling and the dilution tunnel is often referred to as the constant volume sampler (CVS). Since the exhaust flow varies with engine operation but the total flow through the CVS is fixed, the dilution ratio varies during a test.

Heavy-duty dilution systems in Europe tend to be twin stage systems with a small secondary dilution system in series from the main CVS. This secondary dilution system takes a fixed proportion of the flow from the CVS and dilutes it by a preselected ratio. The main aim of this step is to reduce the temperature of the diluted exhaust.

Diluted exhaust is drawn from the secondary dilution system through a filter. The sampled flow rate must be virtually constant and, by definition, proportional to the total flow through the CVS.

In the validation exercise, all the full-flow dilution systems were employed with secondary dilution systems for mass measurements, but particle number measurements were drawn directly from the primary CVS.

The main issues with CVS dilution systems are:

- The transfer time between engine emission and measurement of real time diluted gases in the CVS makes the identification of real time emissions effects complex

- Hydrocarbons and PM are known to deposit and release from the transfer system between the exhaust manifold and CVS

- Finally, the full flow CVS is a large, often ceiling-mounted, tube which is difficult to remove and consequently to clean. Carryover of PM emissions from previous tests may result in high background levels

Full flow dilution systems provided by Horiba and AVL were tested in this work.

#### Partial Flow dilution

Partial flow dilution systems (PFDS) are simpler, more compact and less expensive than CVS.

In a PFDS, a fraction of the raw exhaust (a partial flow) is sampled and diluted. However, the transfer flow from exhaust to tunnel must be proportional to the total flow through the exhaust: In Figure 5; Q1 must constantly change during the transient cycle and this is achieved by varying the flow of dilution air that is added (Q3). As in the CVS, the total flow, Q2, remains constant, but unlike the CVS the entire tunnel flow is drawn through the PM filter.

If an additional flow (Q4) is drawn for further mass or number measurements, an identical increase in the transfer flow, Q1 occurs. This reduces the dilution ratio in the tunnel and would increase the measured PM, so an equivalent flow to Q4 must either be added back upstream of the flow measurement device (which is positioned downstream of the PM filter) or the changes in dilution corrected automatically by the software.

During preliminary work for this programme, JRC performed a comparison of PN measurements from partial flow systems which had both physical correction for the removal of Q4 (i.e. the flow was replaced) and software correction for the removal of Q411. These two approaches were shown to give equivalent results, but during the actual test programme only the software correction approach was employed.

Figure 5: PFDS – Principle of Dilution

To facilitate the changes in exhaust flow rate and dilution flow (Q3), real-time fast flow measurement and rapid changes in flow rate are required.

The main challenge for partial flow dilution systems is maintaining proportionality with the exhaust flow rate. Measurement procedures for particulate emissions using partial flow dilution systems and of gaseous emissions from raw exhaust gases under transient test conditions are defined in an International Standards Organisation publication ISO16183:2002 [].

Partial flow dilution systems provided by Horiba [], AVL [] and Control Sistems [] were tested in this work. A PFDS system is also available from Sierra [].

All these systems are fully compliant with ISO16183:2002 and in principle operate almost identically. The main differences between systems are the methods by which flow is measured and controlled.

## Particulate Mass Measurements

### Full Flow Measurements

In the PMP ILCE_LD the filter-based PMP particulate mass measurement method was employed as the reference method. For conventional Diesels this has been shown to give results consistent with the current regulatory particulate mass measurement method [].

The development philosophy of the PMP particulate mass measurement system was to adapt the readily achievable elements of the mass method proposed for heavy-duty approvals in the US for 2007, along with selected amendments to improve data quality, to create an enhanced European light-duty procedure.  Consequently, the approach developed for light-duty vehicles in the ILCE_LD was considered directly transferrable to heavy-duty engines' full flow dilution system sampling and the main additions to the standard European method are described in the following sections.

Application of highly efficient dilution air filters for particles and hydrocarbons that reduces mass contributions from the dilution air to near zero

The application of a cyclone pre-classifier with a 50% cut-size at between 2.5µm and 10µm to limit the contribution of re-entrained and wear materials to the filter mass

External heating of the filter holder and transfer tubing to permit aerosol stabilisation of >0.2s at 47°C +/-5°C prior to sampling and to ensure close control of the filter face temperature to 47°C +/-5°C. External heating was achieved by either direct surface heating (most labs) or by situating the cyclone, transfer tubing and filter holder in an enclosed vessel. In the second case, the sample probe in the CVS was also heated.

The use of a single 47mm filter rather than primary and back-up filters to eliminate weighing errors and the back-up filter as a source of volatile artefact

The filter medium provides at least 99% filtration efficiency for 0.3µm particles at 35l/min (~50cm/s filter face velocity).

Controlled filter face velocity range (50cm/s to 80cm/s) to improve reproducibility

Definition of PMP Particulate Mass

Despite the changes introduced to the method, the PM definition remains broadly unchanged from that used previously: all materials sampled from a dilution tunnel using the prescribed method on to a single filter at between 42°C and 52°C.

#### Test Facilities

The particulate measurement equipment employed by the participating test labs was constructed to meet the requirements of the ILG_HD. All laboratories used full flow dilution systems equipped with secondary dilution tunnels. The dimensions of these systems, flow rates and residence times were subject to some differences as shown in Table 5.

CVS tunnel residence times were controlled to 1.6s to 2.3s range, but there was a much larger range in secondary tunnel dimensions and residence time (0.4s to 7.8s). As discussed later, these differences did not have a measurable impact on observed PM levels.

Table 5: Principal Differences Between CVS Systems – Test Labs

|  | AVL MTC | JRC | Ricardo | UTAC | EMPA |
| --- | --- | --- | --- | --- | --- |
| CVS flowrate [Nm3/min] | 72 | 80 | 60 | 80 | 80 |
| CVS length [cm] | 500 | 470 | 450 | 575 | 470 |
| CVS diameter [cm] | 50 | 47 | 45 | 45 | 47 |
| CVS Heat exchanger | No | Yes | Yes | Yes | Yes |
| Preclassifier cutpoint [um] | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 |
| Secondary tunnel flowrate [lpm] | 50 | 50 | 60 | 50 | 40 |
| Secondary tunnel DR | 2:1 | 2:1 | 2:1 | 2:1 | 2:1 |
| Secondary tunnel length [cm] | 30 | 64 | 100 | 30 | 200 |
| Secondary tunnel diameter [cm] | 8 | 8.6 | 10 | 8 | 1.4 |

For physical collection of PM, all labs used TX40 filters, but some labs used current PM holders without a back-up and other labs used US07 style holders. Different filter holders did not have a measurable impact on observed PM levels.

### Partial flow Measurements

Partial flow dilution measurements of PM were not undertaken in the ILCE_LD, but development work was undertaken to refine the procedure in the working group that developed the ISO standard [13]. This procedure has many parallels with the PMP full flow method:

- Efficient dilution air filtration

- Filter face temperature control is permissible

- 47mm filters are permitted

- The same filter media are mandated

On these bases, it was considered wise to conduct partial flow testing during the HD_ILCE according to the requirements of ISO16183, but to align sampling parameters where possible with the full flow method. Parameters to be matched were prescribed in the inter-laboratory guide, but these included: filter face velocity, filter medium, filter diameter and dilution air quality.

The definition of PM sampled from a partial flow dilution system: all materials sampled from a dilution tunnel using the prescribed method on to a single filter at 47°C +/-5C.

#### Test Facilities

The particulate measurement equipment employed by the participating test labs was constructed to meet the requirements of ISO16183, with operating parameters specified in the ILG_HD. All laboratories used commercially available PFDS systems. The dimensions of these systems, flow rates and residence times were subject to only minimal differences as shown in Table 6.

Table 6: Principal Differences Between PFDS Systems – Test Labs

|  | AVL MTC | JRC | Ricardo | UTAC | EMPA |
| --- | --- | --- | --- | --- | --- |
| System | Smart Sampler | Smart Sampler & PSS-20 | Horiba Mini Dilution Tunnel | PSS-20 | Smart Sampler |
| PM flowrate [g/s] | 1.08 | 1.08 | 1.205 | 1.08 | 0.86 |
| Split ratio | 0.06% | 0.06% | 0.09% | 0.06% | 0.06% |

As seen with the CVS data, using different filter holders did not result in measurable differences in PM levels.

## Particle Number Measurement Systems

### Principles of the Measurement System

Description and diagram of PMP particle number measurement system: elements and function. Touch upon PCRF and dilution factor-based systems.

### Golden Particle Measurement Systems

Two nominally identical particle number measurement systems were circulated across the participating laboratories for the concurrent determination of the particle number emissions from a full flow CVS tunnel and a partial flow system. These are referred to as Golden Particle Number Systems (GPMS), as they served as an internal standard providing a link between testing at the various laboratories.

The GPMSs selected for this study were two prototype Solid Particle Counting Systems (SPCS) developed by Horiba []. The selection of this particular system was not based on its performance in terms of the criteria specified in the light duty regulations, but rather on the intention to evaluate alternative candidate systems to the GPMS that had been employed in the PMP Light Duty Interlaboratory Correlation Exercise [7].

#### SPCS description

The SPCS unit consists of a hot-diluter (PND1), an evaporation tube (ET), a cold diluter (PND2) and a condensation particle counter (PNC). A flow schematic of the SPCS unit is shown in Figure 6. The aerosol first enters a temperature controlled cabinet where it is diverged into a bypass flow, the sole purpose of which is the decrease of the residence time in the sampling line, and the sample flow. The sample mass flowrate is measured in real time by an orifice flowmeter (CFO), taking into account the temperature and the pressure of the sample as determined with a thermocouple and a pressure transducer, respectively. The sample is then diluted in a temperature controlled mixer (HD) with heated – filter dilution air supplied at an adjustable flowrate by means of a mass flow controller (MFC1).

A small fraction of the diluted aerosol exiting the PND1 passes through an orifice flowmeter and then enters an externally heated evaporation tube (EU) whose wall temperature is controlled in the range of 300 to 400 °C. During the ~0.5 s residence of the aerosol inside the EU, the volatile particles are vaporized to gas phase. Immediately after exiting the EU the thermally treated aerosol enters a mixer (CD) where it is cooled by filtered-dilution air supplied at an adjustable flowrate by means of another mass flow controller (MFC3). The concentration of the aerosol exiting this secondary diluter is then measured in real time in a TSI’s 3010D condensation particle counter (CPC).

The excess flow from the two dilution stages is sampled with a pump. The dilution ratio of the two diluters is kept constant by supplying make-up air in the two excess lines. Two mass flow controllers (MFC2 and MFC4) continuously adjust the make up air to account for small fluctuations of the sample flowrates measured in real time with the two flowmeters.

#### Operating parameters

All labs participated in this study operated the two Golden SPCS units at the same settings. The temperatures at the units were set at:

- Cabinet temperature: 47°C

- Hot dilution air temperature for PND1: 170°C

- Mixer temperature (HD): 170°C

- Evaporation tube (ET): 350°C

In these prototype units, the user has to specify the desired dilution ratio of each diluter as well as the dilution air flowrates and the bypass flow. The values employed in the campaign were:

- Primary dilution ratio (PND1): 10

- Primary dilution air flowrate (MFC1): 11.5 lpm

- Secondary dilution ratio (PND2): 15

- Secondary dilution air flowrate (MFC3): 10.5 lpm

- Bypass flowrate: 2 lpm

These values were selected after preliminary experiments conducted at JRC. The dilution ratios fulfill the specifications laid down in R83 and the resulting sample flowrate (~1.3 lpm) is high enough for pressure fluctuations to have an insignificant effect on the stability of the achieved dilution ratios.

Figure 6: Flow schematic of the prototype SPCS

### Other PMP type systems

A number of alternative candidate systems operating on the pre-conditioning and measurement principles required by PMP have also been used in parallel by the participating labs. However, calibration data demonstrating compliance with PMP performance requirements has not necessarily been provided for all of these systems. These are briefly described below.

#### Nanomet:

This system consists of a primary rotating disk diluter heated at 150°C, a 1m transfer line, an evaporation tube operating at 300 °C and a secondary simple mixer diluter. Two Nanomet systems were tested at JRC (one of which was the GPMS employed in the ILCE_LD) and one at Ricardo. The LD_GPMS system used a TSI 3010D CPC as did the Nanomet employed at Ricardo. The second Nanomet system tested at JRC employed a TSI 3010 CPC modified to replicate the performance of a 3010D (operating at evaporator - condenser temperatures that provided 50% detection efficiency at 23nm) in some tests, and a TSI 3790 CPC in other tests.

#### APC:

This system consists of a primary chopper diluter, a 2m transfer line heated at 150°C, an evaporating tube operating at 350 °C and a secondary dilution stage operating with dilution air at ambient temperature. One APC system was tested at JRC during the second measurement campaign and one at AVL MTC. These systems utilize TSI 3790 CPCs.

#### Dekati dual ejector and evaporating tube system:

In this system the first dilution is applied by a Dekati ejector diluter equipped with a heating mantle (150°C) using heated, conditioned (dehumidified, HEPA filtered) dilution air at 150°C. The diluted sample is then thermally treated in an evaporation tube heated to 330°C and then further diluter in a secondary Dekati ejector using conditioned dilution air at ambient temperatures. This system was employed at JRC during both measurement campaigns, and was tested with three different CPC models: a TSI 3010D, a TSI 3010 and a Grimm 5.404 (the latter two modified in order to shift the 50% detection efficiency to 23 nm).

#### EMPA’s dual ejector and evaporating tube system:

This system consists of a Palas ejector diluter (Palas VKL-10E) heated at 150°C (by means of a heating mantle) operating on conditioned dilution air at 150°C, an evaporation tube operating at 350°C, and a secondary Palas ejector diluter (Palas VKL-10) operating on conditioned air at ambient temperature. Two nominally identical systems were employed at EMPA, one of them sampling from the CVS tunnel (using a TSI 3790 CPC for particle detection) and the other directly from the exhaust (using a TSI 3010 CPC for particle detection).

### PMP like systems:

Some tests were also conducted at JRC using sampling systems employing a Dekati thermodenuder operating at 300°C for the thermal treatment of the aerosol. This replaced the evaporation tube of the PMP type systems. The carbon absorbent section of this particular thermodenuder has an annular design. During the tests, the inner cylinder was cooled by forced convection (using compressed air) and the outer one by natural convection. The thermodenuder operated at a flow rate of 10 lpm which corresponds to a residence time of 0.3 seconds in the heating section and 2.7 seconds in the denuder. The thermodenuder sampled from the CVS tunnel either directly or via a heated (heating blanket at 150°C) Dekati ejector operating with heated dilution air at 150°C. A TSI 3790 CPC was employed for particle detection.

### Additional Instrumentation

In addition to the aforementioned PMP type systems, the participating laboratories have also employed additional aerosol instrumentation in order to get a better insight into the characteristics of the emitted particles. These included:

EEPS:

A 3090 Engine Exhaust Particle Sizer Spectrometer (EEPS) (TSI Inc.) was used during a limited number of tests. EEPS measured particle size distributions with a maximum data rate of 10 size distributions per second (although averages over 1 second were used in the graphs of this study). It measured particle sizes from 5.6 to 560 nm with a sizing resolution of 16 channels per decade (a total of 32 channels). At the instrument’s inlet there was a cyclone with a 50% cut-size at 1 μm (inlet flow rate 10 lpm).

Dekati Mass Monitor:

A Dekati Mass Monitor (DMM) was employed at JRC during some tests conducted after the second measurement campaign. The DMM sampled from the CVS tunnel via a Dekati thermodenuder operating at 300°C, and was used to measure the mass concentration of non-volatile particles in real time.

AVL Soot Sensor:

An AVL’s 483 soot sensor was employed during some of the formal tests conducted during the second measurement phase at JRC, measuring the mass of soot in real time. The soot sensor sampled directly from raw exhaust at a constant dilution ratio of about 2.

TSI SMPS:

A TSI’s 3936L SMPS (consisting of a TSI 3080L DMA and a TSI 3010 CPC) was employed during some preliminary tests at JRC operating at sample/sheath flowrates of 0.9 and 9 lpm respectively.

TSI 3025A CPC:

During some of the preliminary tests conducted at JRC two TSI 3025A CPC units having a 50% counting efficiency at 3 nm were employed. One of them was used to sample directly from the CVS tunnel, and the other to sample from the GPMS connected to the CVS tunnel, and in parallel to the Golden CPC (50% efficiency at 23nm). These tests allowed for the determination of the number concentrations of volatile and non-volatile particles in the 3-23 nm size range. This provided the means to investigate whether significant emissions of non-volatile particles smaller than 23 nm, suggested by recent studies [], are also observed with the Golden engine.

## Test Programme

### Participating Laboratories

Four laboratories in EC member states, and one in Switzerland, were participants in the test programme. The test labs, timing and final test order are given in Table 7. JRC also conducted additional experiments prior to and following the formal testing, these experiments are not included in the assessment of measurement repeatability and reproducibility.. Preliminary experiments, undertaken to refine the inter-laboratory guide are discussed in the next section.

Table 7: Test Laboratories and Timeline

| Dates | Test Laboratory | Location | Testing |
| --- | --- | --- | --- |
| Jan – Feb 2008 | JRC | Ispra, Italy | Preliminary Experiments |
| Mar-Apr 2008 | AVL-MTC | Sweden | Formal Testing Lab#1 |
| May – Jun 2008 | JRC | Ispra, Italy | Formal Testing Lab#2 |
| Dec 2008 – Jan 2009 | Ricardo | UK | Formal Testing Lab#3 |
| Feb - Apr 2009 | UTAC | France | Formal Testing Lab#4 |
| Apr – Aug 2009 | EMPA | Switzerland | Formal Testing Lab#5 |
| Aug– Oct 2009 | JRC | Ispra, Italy | Formal Testing Lab#2rpt |
| Nov-09 | JRC | Ispra, Italy | Additional Experiments |

### Summary of Preliminary Experiments

A number of experiments were conducted in order to better define the measurement approaches described in the inter-laboratory guide. These were based upon concerns regarding the possible differences in emissions between the light duty vehicles that the measurement procedures were developed for, and the heavy-duty engines to be the subject of the next phase of PMP work. Further details can be found elsewhere [10]. In particular, the differences in PM chemistry, the aftertreatment devices to be used, characteristics of different dilution systems and possible differences in the nature of particles: their sizes and origins, were of concern. These experiments can be classified into four groups as follows:

- Expt 1 - Background and filter tests

These experiments considered the CVS and partial flow systems’ background levels for mass and the effect of the filter medium and filter face velocity (ffv) on the PM emissions

- Expt 2 - Sampling parameters

These experiments considered the impact and necessity of using a pre-cyclone with the Horiba SPCS systems. In addition, the two SPCS systems supplied by Horiba had different transfer lines lengths (1m insulated, 4m heated). The impact of these lines on the PN results was determined.

- Expt 3 - Pre-conditioning and continuity protocol

In the ILCE_LD, it was determined that to improve the repeatability of particle number results a purge of pre-existing particles from the exhaust system and a standardised DPF fill-state was required. Experiments were undertaken to determine the minimum required pre-conditioning for the exhaust and after treatment for repeatable measurements.

- Expt 4 - Real time PN emissions

Concerns have been raised in the US that high levels of solid particles may be present, in the exhaust from HD Diesel engines, in the size range below the PMP cut-off (d50) of ~23nm. Measurements were undertaken to determine the presence and magnitude of <23nm solid and <23nm volatile particles from the golden engine.

All experiments were conducted on the golden engine and emissions control system and using the fuel and lubricant described in Section 3.5.

#### Experiment 1 - Background and Sampling Parameters for PM

A schematic of the sampling system used in these experiments is shown in Figure 7.

Evaluations of filter media and face velocity effects were undertaken during repeat WHSC testing. In all cases tests were conducted using a primary and back-up filter to enable the magnitude of volatile ‘slippage’ from the primary filter to the secondary filter to be quantified. The following measurements were made from the CVS, with results shown in Figure 8.

(a)

- One background (BG) PM measurement followed by 3 hot WHSC repeats

- 70mm TX40 filters, PM sample flow of 60l/lmin (ffv = 43cm/s)

(b)

- One background PM measurement followed by 3 hot WHSC repeats

- 47mm TX40 filters, PM sample flow of 60l/lmin (ffv = 103cm/s)

(c)

- One background PM measurement followed by 3 hot WHSC repeats

- 47mm TX40 filters, PM sample flow of 40l/lmin (ffv = 69cm/s)

(d)

- One background PM measurement followed by 1 hot WHSC

- 47mm TX40 filters, PM sample flow of 40l/lmin (ffv = 69cm/s)

- PM filters baked in an oven at 47°C for 3 hours to remove residual volatiles

(e)

- One background PM measurement followed by 1 hot WHSC

- 47mm Teflon membrane filters, PM sample flow of 40l/lmin (ffv = 69cm/s)

Figure 7: PM Sampling Set-up for Experiment 1

Figure 8: Experiment 1 - Background and Sampling Parameters for PM

The conclusions of these experiments were:

- PM emissions were slightly higher than background levels

- Highest PM mass was collected on 70mm filters with a sample flow of 60l/min  (ffv of ~40cm/s)

- Similar masses were collected on 47mm filters at ffvs of ~70 to 100cm/s (40 to 60l/min)

- Baking TX40 filters has no beneficial effect

- Teflo filters appear to collect less mass than 47mm filters

- Secondary (backup filters) collect ~ 30% of primary filter mass from sample filters

Contributions to test protocol

- No substantive changes to light-duty PM protocol

- Employ 50 l/min flow rate for PMP tests (for full flow and partial flow systems with 47 mm TX40 filters).

#### Experiment 2 – PN Sampling Parameters

In these experiments 2 SPCS systems were used in parallel during various periods of transient engine operation.

(f)  Transfer line effects (Figure 9)

- One SPCS sampling directly from the CVS with a 1m insulated line

- One SPCS sampling directly from the CVS with a 4m heated line

(g) Cyclone effects (Figure 10)

- One SPCS sampling directly from the CVS with a 1m insulated line

- One SPCS with 4m heated line sampling from the CVS via a cyclone operating with a 4(m cut

Figure 9: Minimal Impact of Heated Line on PN Results

Figure 10: Minimal Impact of Cyclone on PN

The conclusions of these experiments were:

- The differences in PN emissions measured with a 4 m heated line (at 47°C, with 1.8s residence time) and a 1m insulated line (0.5s residence time) were minimal (<5%). This means that one SPCS can be connected to the CVS through 4 m line (heated at 47°C), while the other SPCS can be connected to the partial flow system with a short (insulated) line without adjustment of results for transport losses.

- The cyclone had a negligible effect on the particle number emissions, so for number measurements it is not necessary. However, it is recommended to use one in order to protect instruments’ primary diluters from contamination through deposition of larger particles. Insulation of the cyclone and transfer tubes is required to limit thermophoretic losses during high temperature operation.

Contributions to Test Protocol

- Cyclone mandatory for CVS sampling

- Cyclone optional for partial flow sampling

- Insulation of cyclone and external sampling system to limit thermophoretic losses

- Dedicated SPCS (Serial no. 19) for full flow sampling (with heated line)

- Dedicated SPCS (Serial no. 20) for partial flow sampling (with 1m insulated line)

#### Experiment 3 – Preconditioning Protocols

A daily preconditioning protocol was required that first used a high exhaust temperature steady state (~600°C) to passively regenerate the DPF, and then used a lower temperature non-regenerating condition to add a standardised quantity of soot to the DPF. This protocol was used at the end of each day to re-baseline the loading state of the DPF prior to the next day’s run-through of the test matrix. To avoid very long test days, it was desirable that the entire process was shorter than 2 hours.

- 15 minutes at mode 10 was determined as the suitable engine operation and minimum time required to passively regenerate the DPF, eliminating the stored soot. This was determined by running ESC 10 and monitoring how long it took for post DPF solid particle emissions to stabilise (Figure 11, green line, 400s to 900s).

Mode 10 operation for 15 minutes showed particle number levels consistent with those seen during the 2 hours of mode 10 running used to condition the lubricating oil, thus indicating that the DPF was indeed ‘emptied’ by the 15 minutes period (Figure 11, red and blue lines).

- ESC Mode 7 has a stabilised exhaust temperature of well below 300°C, so passive regeneration at this temperature, even with an oxidation catalyst, is minimal. 30 minutes of operation was chosen to add soot to the DPF  (Figure 11, green line, 1500s to 3300s)

The daily test matrix comprised both cold and hot start tests and it was desirable that each lab tested the cycles with the exhaust and emissions control system at the comparable temperatures. This would reduce variability due to pre-conditioning effects and allow a better assessment of the repeatability of gaseous and particle emissions measurement procedures. To ensure this, cycles without regulatory defined warm-ups were preceded by the continuity protocol.

The hot WHTC was preceded only by the cold WHTC and a 10 minute soak while the WHSC, was preceded by the hot WHTC and the mandatory 10 mins at WHSC Mode 9. No continuity protocols were required for these cycles.

The continuity protocol comprised 5 min at Mode 7 of the ESC and 3 min at idle. Mode 7 was selected after repeat size distribution measurements at this condition showed no evolution in magnitude or size – indicating stability of both volatile and solid particles (Figure 12). By comparison a slightly hotter steady state, ESC Mode 4 showed obvious evolution, related to some passive regeneration.

Idle followed the Mode 7 operation to permit preparation of the exhaust emissions analysers and to limit emissions and fuel consumption without stopping the engine.

Figure 11: DPF Preconditioning –Regeneration (Mode 10) and Fill (Mode 7)

Figure 12: Stable Mode 7 Selected for the Continuity Protocol

Contributions to test Protocol

15 mins ESC Mode 10 plus 30 mins ESC Mode 7 for the daily preconditioning

5 mins ESC Mode 7 plus 3 mins at idle for the continuity protocol

#### Experiment 4 – Real-time PN Emissions

Concerns have been raised in the US that high levels of solid particles may be present, in the exhaust from HD Diesel engines, in the size range below the PMP cut-off (d50) of ~23nm. Measurements were undertaken to determine the presence and magnitude of <23nm solid and <23nm volatile particles from the golden engine.

A variety of steady state and transient emissions tests were performed measuring solid particles from the CVS with an SPCS system equipped with both a TSI 3010D CPC (measuring particles >23nm) and a TSI 3025A CPC (measuring particles >3nm). The difference between the results of these two particle counters indicated the presence of solid particles in the size range 3nm to 23nm. In addition, the same 3025A CPC was used to measure particles directly from the CVS. This permitted the number of volatile particles smaller than 23nm to be determined. Results are summarised in Figure 13.

Figure 13: Particle Number Emissions <23nm

The implications of this figure can be summarised as follows:

For the cycles examined, the concentration of total particles >3 nm measured from the SPCS was generally 50-95% higher than the >23nm non-volatile particles. The difference appears to be higher over the hot WHTC cycle, measuring particle concentrations direct from the CVS (300%). CVS measurements include both solid and volatile particles and this result is therefore likely to be indicative of a higher proportion of volatile particles over this cycle

For the cycles examined the concentration of the non-volatile particles <23 nm was 15-45% higher than the non-volatile particles >23 nm (85% for the cold WHTC).

Impact of Experimental Result on the Test Protocol for the Validation Exercise

While there was some evidence that solid particles <23nm were present, the levels seen were not consistent with the orders of magnitude increases relative to >23nm particles reported from US engines. On this basis, it was considered reasonable to retain the size and volatility range of particles measured in the ILCE_LD for measurements from heavy-duty engines.

### Daily Protocol

Following the completion of the preliminary experiments, the test protocol for the inter-laboratory exercise was finalised and used to update the ILG_HD. The baseline test matrix comprised at least 8 repeats of each of the following tests:

- Cold WHTC

- Hot WHTC

- WHSC

- ETC

- ESC

Test order followed the defined matrix (Table 8), with preconditioning for each cycle set as the regulatory requirement, or the continuity protocol if no regulatory requirement exists. The continuity protocol was defined as 5 minutes operation at ESC Mode 7 plus 3 minutes at idle (as described in experiment 3 of Section 3.10.2).

Table 8: Matrix for Emissions Testing

# STATISTICAL ANALYSES

The variability of the results collected in the PMP Heavy Duty Validation Exercise was quantified using the “random effects analysis of variance” model []. This analysis provides the means for a separate quantification of the repeatability and reproducibility of the measurements.

On the other hand, the equivalency between the results (PM and particle number) obtained from the CVS tunnel, the partial flow system and the alternative particle number systems employed, was investigated by means of calculating the average value and the standard deviation of the percentage differences of each individual result. It has been decided not to employ paired t-tests for this type of checks as the particular methodology is prone to identifying statistically significant differences when in fact the result are practically equivalent [], with the ability to discriminate between statistically and practically significant differences strongly affected by the sample size.

## Definitions

### Basic statistical concepts

Before describing the statistical analysis used for the evaluation of the PMP methodology it is important to distinguish between the statistical concepts of accuracy, precision and trueness.

Accuracy represents the degree of agreement between the results obtained from a test method and the true or ‘accepted’ true value. On the other hand, precision refers to how closely the independent measurements agree with each other, while trueness refers to the closeness of agreement between the arithmetic mean of a large number of test results and the “accepted” reference value. In most cases, however, the true value is unknown and therefore only precision statements can be developed. Precision is a qualitative concept which can be expressed numerically only in terms of its opposite, that is the variance or standard deviation.

The variance in the results obtained from a test method is due to some random variations of the properties being measured but also due to the fluctuation of some factors affecting the outcome of the test method. These factors are generally the equipment used, the calibration of the equipment, the operators using the equipment and the environmental variables.

When the test method is performed in one laboratory in the shortest practical period of time, by the same operators, using the same equipment on – ideally – materials taken from a single quantity of homogeneous material, then the aforementioned factors remain reasonably constant and the variance in the results is referred to as the within laboratory (intra-laboratory) variability (σ2). However, when the test method is performed at different laboratories these factors vary considerably, leading to even greater variability. The variability induced because of performing the tests in different physical environments is the between laboratories (inter-laboratory) variability.

The results obtained from one laboratory are said to be satisfactory if they are both repeatable and reproducible. Repeatability is ensured when the above mentioned factors remain reasonably constant. This would be reflected where inter-laboratory variance was similar to intra-laboratory variance. Furthermore, the results obtained from one laboratory are said to be reproducible if the divergence of their mean value is not significantly greater than the variability range which would be expected on the basis of the inter-laboratory variability.

### Intra- And Inter-Laboratory Variance Estimators

The random effects analysis of variance model provides the means for the quantification of the intra-laboratory and inter-laboratories variances σ2 and στ2, respectively. If yij represents the jth result obtained from the ith laboratory participating in an interlaboratory correlation exercise (also known as round robin) conducted at p different laboratories, and ni is the number of results provided from the ith laboratory then the estimators of σ and στ2 are [23]:

Where:

,  and

### Statistical outliers

One of the basic assumptions of the analysis of variance model is that the results obtained from each laboratory are equally variable (an situation known as homogeneity of variance). If the variability in one particular laboratory is significantly different from the rest of the laboratories the particular laboratory is said to have repeatability problems and can be characterized as an outlier. The ASTM 691-99 and the ISO 5725-2:1994 standards provide the means for identifying the laboratories having repeatability problems, by utilizing the repeatability index (k). The k statistic for the results obtained from a laboratory z is defined as

where sz represents the standard deviation of the results obtained from laboratory z, while sr is the estimated intra-laboratory standard deviation (square root of sr2).

Furthermore, the two standards suggest the use of the reproducibility index (h) in order to identify laboratories measuring significantly higher or lower results. The h statistic for the results obtained from laboratory z is defined as:

where  represents the average value of the results obtained from laboratory z, μ is the pooled mean value from all labs and s0 represents the range where the difference  is expected to vary due to the intra and inter-laboratory variabilities (sr2 and sL2, respectively).

The statistical analysis underlying the definition of these two statistics is the hypothesis testing of two variances and two means respectively, and therefore some critical h* and k* values can be derived by assuming a type-I error (a). Any h or k value greater or equal to the corresponding critical h* or k* is indicative of reproducibility or repeatability problems, respectively. There are also additional patterns indicating problems like one lab having positive (or negative) h values and all the rest negative (or positive), or one lab having too high or too low k values for all tests compared to the rest of the labs.

The two aforementioned standards provide the equations and critical values for the case of balanced samples (equal number of results from each lab). However, in this exercise not all labs performed exactly 8 repetitions of the test protocol. Additionally, for various technical reasons, some of the data were excluded and therefore the final sample ended up statistically unbalanced. The following equations apply in the more general case of unbalanced samples []:

where t is the α/2 percentage point of the Student’s t-distribution with p-2 degrees of freedom, and F the α percentage point of the F-distribution for nz-1 degrees of freedom for the numerator and [N-p-(nz-1)] degrees of freedom for the denominator.

Following the recommendations of the ISO standard, the critical h and k statistics were calculated at a significance level (α) of 1 %.

The standard deviations contain information on the absolute level of each property tested. In order to compare the variability of properties differing by orders of magnitude it is necessary to normalize the standard deviation with respect to the average emission levels. For this purpose, the within-laboratory (sr/μ) and between variability (sL/σ) coefficients of variation (CoV) are used instead in this document. For convenience these quantities are also referred to as within and between laboratories variability.

# CALIBRATION AND VALIDATION

## Mass systems

Filter mass measurement procedures in the weighing room (or chamber) were conducted according to the requirements of the ILG_HD. During the entire exercise, no substantive issues were reported with the weighing procedures of the HD_ILG at any of the test labs.

Figure 14 shows the results of weighing environment validation exercises undertaken at Ricardo. During a calendar month there are typically no deviations from temperature range (19°C to 25°C), relative humidity range (37% to 53%) or in the performance of the balance (50.002µg ± 5µg). However, several borderline results from reference filter weighings suggests that variability (of 47mm TX40 filters) increases such that ± 10µg range could be required instead of the ± 5µg used in the R83. This is in line with the requirements of the WHDC.

Figure 14: Weighing Room Validation Exercises - Ricardo

## Number systems

Prior to the commencement of the ILCE_HD, the two SPCS systems and their PNCs (3010D) were calibrated by their manufacturers. These calibrations were performed prior to the finalisation of the PMP calibration methodologies, including that for particle concentration reduction factor (PCRF). The calibrations performed are outlined in the following sections.

### SPCS Calibration

The two SPCS systems used as Golden Instruments during the ILCE_HD were supplied by Horiba and calibrated according to their own internal procedures.

SPCS-20 was subjected to a comprehensive characterisation, while SPCS-19, built in parallel with identical componentry, was subjected to a dilution ratio gas calibration only but compared with the SPCS20 within the main PMP programme and shown to be highly similar.

The calibration approach used for the SPCS-20 has been published by Horiba [] and is briefly outlined below. This included determination of VPR penetration using polydisperse aerosol and monodisperse aerosols and dilution ratio validation using a propane span gas and a flame ionisation detector. The removal efficiency of the VPR for 30nm tetracontane (n-C40) particles was determined according to the procedure previously described by Horiba [].

#### Removal Efficiency for Tetracontane Particles

The apparatus used by Horiba to determine the volatile particle removal efficiency (RE) of SPCS-20 is shown in Figure 15. The C40 aerosol generator heats up to the boiling point of tetracontane and the neutraliser and DMA are used to select monodisperse 30nm particles. The concentration of these particles is reduced by a room temperature ejector diluter to below 10,000/cm3 and they then enter the evaporation unit (EU) at 300°C. Upstream and post-EU concentrations are compared to evaluate the RE

Figure 15: Apparatus for the Evaluation of Volatile Particle RE

A typical result for this particle removal test, showing efficiency of 99.99% for 30nm n-C40 particles is given in Figure 16.

Figure 16: VPR RE of 99.99% for 30nm n-C40 particles

#### Solid Particle Penetration Using Polydisperse Aerosol

The apparatus used by Horiba to determine the solid particle penetration of SPCS-20 using polydisperse NaCl is shown in Figure 17. Temperatures of the PND1 dilution air, mixer, EU, and sample flow were controlled to 150°C, 150°C, 320°C, and 47°C, respectively. These temperatures were used for normal SPCS operation at the time of these experiments.

Sodium Chloride (NaCl) aerosol is generated with the atomizer and residual moisture in the aerosol is then removed with an efficient diffusion dryer. A by-pass is placed upstream of the diffusion dryer to vent excess aerosol flow. In the neutralizer, the aerosol is charged to Boltzmann equilibrium. An ejector diluter is used to provide an aerosol at the correct concentration level and to enable the flow to be controlled by the SMPS. The concentration and the size distribution then remain constant while the upstream (raw) and the downstream (diluted) size distributions of the VPR are measured. Comparisons of these distributions for specific size ranges (Figure 18) and at different dilution ratios provide the penetration efficiencies (Figure 19).

At the typical dilution factor used in the PMP work (~150), the polydisperse penetration was >98%.

Figure 17: Apparatus for Evaluation of Solid Particle Penetration (Polydisperse)

Figure 18: Size Distributions (downstream dilution correction of 1050 applied)

Figure 19: Penetration at Different Dilution Ratios (Polydisperse)

#### Solid Particle Penetration Using Monodisperse Aerosol

The apparatus used by Horiba to determine the solid particle penetration of SPCS-20 using monodisperse NaCl is shown in Figure 20. Temperatures of the PND1 dilution air, mixer, EU, and sample flow were controlled to 150°C, 150°C, 320°C, and 47°C, respectively. These temperatures were used for normal SPCS operation at the time of these experiments.

Sodium Chloride (NaCl) aerosol is generated with the atomizer and residual moisture in the aerosol removed with an efficient diffusion dryer. A by-pass is placed upstream of the diffusion dryer to vent excess aerosol flow. In the neutralizer, the aerosol is charged to Boltzmann equilibrium. An ejector diluter is used to provide an aerosol at the correct particle concentration level (less than 10,000/cm3 upstream of the VPR for each monodisperse particle size) and to enable the flow to be controlled by the CPC. The DMA is employed to select particles of 30nm, 50nm and 100nm electrical mobility diameter.

Pre and post VPR concentrations are compared to determine penetration levels. Results of these comparisons at several dilution ratios are shown in Figure 21, Figure 22 and Figure 23 for 30nm, 50nm and 100nm particles respectively.

Figure 20: Apparatus for Evaluation of Solid Particle Penetration (Monodisperse)

Figure 21: SPCS Penetration of 30nm Particles

Figure 22: SPCS Penetration of 50nm Particles

Figure 23: SPCS Penetration of 100nm Particles

These data convert to a range of PCRF values from 1.06 (taking all the highest results irrespective of dilution ratio) to 1.20 (all the lowest penetrations). This compares reasonably well with the value of 1.25 obtained at JRC (Error! Reference source not found.) using the finalised PCRF procedure.

#### Dilution Ratio Calibration Using Propane Span Gas

Comparisons of HC concentrations, corrected for ambient levels, were made upstream and downstream of the VPR at a range of dilution ratios, using a flame ionisation detector. One series of experiments was performed with a fixed primary dilution ratio and another with a fixed secondary dilution ratio. Results are shown in Figure 24 and Figure 25.

Results showed that dilution ratio settings were within ±4% of the nominal value.

Figure 24: Dilution Ratio Calibration – Fixed Secondary Dilution

Figure 25: Dilution Ratio Calibration – Fixed Primary Dilution

### 3010D Calibration

Particle number counters used with the two SPCS systems were calibrated according to the procedures described by TSI []. This uses the primary method in which particle counting is verified by comparison with the response of a electrometer. The counting efficiencies of the PNCs were determined through the use of an electrospray generated polyalpholefin aerosol classified in a DMA.

Certificates of conformity for the two PNCs used with the SPCS systems are given in Appendix 2. These data were supplied directly by TSI [].

### Validation Exercises

No substantive issues were observed with the PN validation procedures of the HD_ILG at any of the test labs during the programme. These included zero checks of the PNC and SPCS systems and flow checks of the PNC. Functional checks (indicative monitoring of flows and temperatures) were undertaken on a continuous basis and any issues resolved via routine maintenance.

# EMISSIONS RESULTS

## Full Flow and Partial Flow PM

Phase II of PMP made recommendations for introducing a solid particle number measurement procedure and also for improving the current EU and ECE regulatory particulate mass measurement procedure. The aim of Phase III of PMP was to validate the measurement techniques recommended in Phase II. The ILCE_LD validated both measurement techniques for use in light duty vehicle emissions measurement. The ILCE_HD also included measurements to evaluate the revised particulate mass measurement technique recommended by PMP Phase II when applied to heavy duty engine emissions measurement. However, international agreement on improvements to particulate mass measurement techniques for heavy duty engine emissions testing has already been reached within Global Technical Regulation No.4. Many of the particulate mass measurement improvements adopted in GTR No.4 are consistent with the PMP Phase II recommendations. For these reasons it is not the intention of the PMP informal group to propose amendments to regulatory, heavy duty particulate mass measurement procedures. However the results and conclusions of the measurements made using the revised particulate mass measurement procedures during the ILCE_HD are reported here for completeness.

Particulate mass data from full and partial flow dilution systems are both discussed in this section.

### PM - Repeatability

Repeatability levels for all labs are expressed as single CoV values that express overall intra-lab variability for each emissions cycle (see Section 4.1.2).

Figure 26 shows the repeatability of the 5 test matrix cycles and the composite weighted WHTC result for the CVS-based PM method.

Three results are shown for each cycle, and these include:

- PMCVS1: All data from all labs (excepting tests excluded for technical reasons)

- PMCVS2: Outlier analysis iteration 1

- PMCVS3: Outlier analysis iteration 2

Outlier analyses found no exclusions.

The best repeatability of ~34% was seen from the cold start WHTC cycle, with all other cycles showing between 50% and 56%. Filters from emissions tests revealed that the cold start WHTC showed both the highest sample masses and visible grey staining indicating the presence of some elemental carbon. Other cycles’ filters did not show the same discoloration.

Figure 26: Repeatability of the CVS PM Method

Figure 27 shows the repeatability of the 5 test matrix cycles and the composite weighted WHTC result for the PFDS-based PM method.

Three results are shown for each cycle, and these include:

- PMPFDS1: All data from all labs (excepting tests excluded for technical reasons)

- PMPFDS2: Outlier analysis iteration 1

- PMPFDS3: Outlier analysis iteration 2

Outlier analyses led to the exclusion of all PM results from UTAC which showed PM levels significantly higher than PFDS results from any other lab. Interestingly, this lab’s PFDS results showed very similar levels to its CVS-measured results. Following this analysis, the resulting repeatability levels were typically 20% to 30% across all emissions cycles.

Figure 27: Repeatability of the PFDS PM Method

### PM - Reproducibility

Reproducibilities are given as single CoV values that express overall inter-lab variability for each emissions cycle (see Section 4). Reproducibility levels for the CVS and PFDS PM methods are shown in Figure 28. Data generated following two rounds of outlier iterations are shown.

CVS PM reproducibility levels were typically in the range 35% to 55%, averaging 42.7% for the 5 emissions cycles in the test matrix. PFDS PM reproducibility levels ranged from ~30% to ~45%, averaging 36.1%.

The lower CoVs from the PFDS systems probably reflect the greater consistency of tunnel background levels in the partial flow system compared to the CVS.

Figure 28: Reproducibility of PM Methods

### PM – Filter Weights, Background Levels and Background Subtraction

Testing at Ricardo included both CVS and PFDS PM measurements, and a comparison has been made of the filter masses collected from background and cycle testing.

#### Filter Weights

In all cases filter masses proved to be higher from the CVS than from the partial flow system (Figure 29). The greatest differential between measurement systems proved to be from the WHTC cycles, with results between the systems closest from the ESC cycle.

Filter loadings seldom exceeded 50µg with partial flow sampling (collected masses ranged from a high of 59µg (ESC) to a low of zero (also ESC)). CVS levels were both higher and covered a wider range: from 346µg over one cold WHTC down to 31µg (ESC).

Figure 29: Sampled Filter Masses – Various Cycles, Both Measurement Systems

Figure 30 puts measured filter masses into context of the tunnel background filter levels recorded:

- It is clear that from the CVS, the highest sample mass is higher than the highest tunnel background, but the lowest sample mass is lower than the lowest tunnel background.

- From the partial flow system, the highest and lowest sample masses are roughly equivalent to the highest and lowest tunnel background masses.

The observations from this testing were as follows:

- Mass emissions measured by the partial flow system may be indistinguishable from the dilution tunnel background. In which case, mass emissions from all cycles are effectively zero.

- Mass emissions measured by the CVS system were in some cases higher than the tunnel background levels, but it is also possible that the CVS tunnel background levels at this laboratory may be unrepresentative.

If the first observation and comment (all mass emissions measured by the partial flow system were effectively zero) is correct, and the partial flow results are valid, then the highest CVS results must have been seeing a greater tunnel background contribution during the test than from the pre-test measurement.

Figure 30: Filter Masses - Samples and Tunnel Backgrounds Compared

This testing suggested that tunnel background subtraction of filter masses from partial flow and a substantial number of full flow tests is likely to give net PM results of 0 mg/km. This may be a true indication of the mass emissions of this engine with DPF when the resolution of the mass method is considered.

Tunnel background PM was also frequently measured at EMPA and JRC.

Tunnel background PM filters were drawn prior to the cold start WHTC emissions test each day. The mass on each tunnel background filter was then treated as a cycle emissions result (masses were adjusted for differences in sample times between cycles) and the corresponding mg/kWh emissions figure calculated using engine data from that day's emissions tests.

#### CVS PM Tunnel Background

As Figure 31 shows for CVS-sampled PM data from both EMPA and JRC, only measurements from ESC cycles were above tunnel background levels. The ESC cycle has a substantial period of operation at high exhaust temperatures and this may lead to emissions of low volatility compounds that are efficiently collected and then retained by the filter. Filters from other cycles collect higher volatility materials from the exhaust and dilution air, but these can be released following acquisition through volatilisation or through a washing effect as further aerosol is drawn through the filter.

It is worth noting that except for hot start WHTC measurements at EMPA, both samples and  during the ILCE_HD results were above the limits of detection (3 standard deviations of the blank measurement) for the various cycles.

Figure 31: Tunnel Background and Sample PM levels - CVS

These data suggest that the CVS PM method is capable of resolving PM emissions from ESC tests from tunnel background levels. Results from other cycles, including the cold WHTC are subject to high uncertainty and would reduce to zero if tunnel background subtraction was undertaken.

#### PFDS PM Tunnel background

Tunnel background and sample filter comparisons were also made from partial flow dilution systems at JRC and EMPA (Figure 32).

EMPA results reflected the CVS results, where all cycles’ data except ESC were similar to the tunnel background levels.

JRC results, conversely, showed that it is possible to discriminate PM samples from the tunnel background, but this discrimination is poorest from the hot and cold WHTC cycles.

Tunnel background correction of the JRC PFDS results would reduce the emissions levels from cold start WHTC to (generally) <1mg/kWh, hot start WHTC to 0.5g/kWh or less, WHSC to ~1.2g/kWh, ETC to 0.5g/kWh or less and ESC to <2mg/kWh.

Figure 32: Background and Sample PM levels - PFDS

#### Limits of Detection (LOD) for Mass Methods –CVS and PFDS

LOD from tests at EMPA and JRC show that for the hot-start WHTC at EMPA only, the limit of detection was above the Euro V limit (10mg/kWh) for the weighted WHTC cycle. However LOD for partial flow systems were always below 2mg/kWh.

Table 9: Limits of Detection – PM Methods at JRC and EMPA (mg/kWh)

|  | EMPA CVS | EMPA PFDS | JRC CVS | JRC PFDS |
| --- | --- | --- | --- | --- |
| C_WHTC | 4.9 | 1.2 | 0.18 | 0.10 |
| H_WHTC | 13.8 | 1.7 | 0.24 | 0.15 |
| WHSC | 6.4 | 0.9 | 0.11 | 0.07 |
| ETC | 1.7 | 1.0 | 0.07 | 0.05 |
| ESC | 2.2 | 0.2 | 0.03 | 0.03 |

### PM – Emissions Levels

Figure 33 (CVS) and Figure 34 (PFDS) show the maximum to minimum ranges of Particulate Matter emissions levels seen from each emissions cycle at each laboratory. Ranges are shown as error bars above and below the mean values. These data are not corrected for tunnel backgrounds.

#### PM emissions from CVS Systems

Emissions levels from the CVS (Figure 33) showed the largest ranges from Ricardo and EMPA, where tunnel background levels were substantially higher than other laboratories. These laboratories also showed some of the lowest emissions values as well as some of the highest, but were sampled according to the prescribed protocols using compliant equipment and thus were not be eliminated as outliers by simple statistical techniques.

The contribution of high and variable tunnel background PM to the results from EMPA and Ricardo is believed to be responsible for the variable PM results from these two labs and this may be related to the recent test history of the facility. In particular, Ricardo had undertaken testing on high bio-content fuels, active regeneration strategies for DPF regeneration and undertaken substantial amounts of non-DPF testing on low NOx calibration engines. All these types of testing would be expected to contribute substantially to both volatile and carbonaceous CVS tunnel backgrounds.

The test protocols were designed to help purge the CVS system of residual tunnel backgrounds, but it is clear that in some cases extreme measures may be required to eliminate historical PM from full flow dilution systems. For this reason it may be necessary to permit the subtraction of a tunnel, rather than dilution air, tunnel background for regulatory PM purposes.

Tunnel background levels of PM in other laboratories were very low (typically <1mg/kWh).

Generally speaking, and excepting some results from Ricardo and EMPA which were higher, PM emissions from all cycles were <6mg/kWh, with no obvious difference in emissions between the cold and hot start WHTC cycles. These levels are substantially below the 10mg/kWh limits set for the weighted WHTC and WHSC at Euro V and expected for Euro VI.

Tunnel background correction of the PM results from Ricardo [] (which showed the highest emissions of all labs) brought them in line with other labs: reducing Cold WHTC results to ~8mg/kWh, hot WHTC results to ~7mg/kWh, WHSC to ~4mg/kWh, ETC to ~1mg/kWh and ESC to <1mg/kWh.

#### PM Emissions from PFDS Systems

PM emissions measured by PDTs (Figure 34) showed narrower ranges than those measured from CVS systems.

Statistical analyses identified PM results from UTAC as systematically higher, with mean values from all cycles at between 4mg/kWh and 7mg/kWh. It is possible that the PFDS used by UTAC had been employed for non-DPF testing and was contributing particles during emissions tests. This is analogous to the high tunnel background PM levels observed by Ricardo from their CVS.

By contrast, PFDS emissions levels from other laboratories, and considering all emissions cycles, rarely exceeded 4mg/kWh.

As seen in the CVS PM data, the emissions levels from cold and hot start WHTC do not appear to differ.

Comparisons between CVS and PFDS mass data are discussed in Section Error! Reference source not found., but emissions from the same engine measured simultaneously from CVS and PFDS appear, in general, to be lower from the PFDS.  This is likely to be related to the fact that almost all the PDTs tested were relatively new and one, tested at Ricardo, had never been used before. Newer dilution systems are less likely to have been exposed to old technology higher PM engines. Tunnel background contributions to PM are therefore likely to be low.

It is widely assumed that dilution systems reach a deposition and entrainment equilibrium where losses to the dilution tunnel walls are balanced by resuspension. This follows a period where particle deposition to tunnel walls is favoured. ‘Clean’ systems may be within the deposition-dominated phase and this may explain the directionally lower PM results seen with the PFDS data from Ricardo.

Figure 33: Maximum and Minimum Ranges of Emissions – CVS PM

Figure 34: Maximum and Minimum Ranges of Emissions – PFDS PM

### PM – Daily Trends

The trends in PM emissions across the day’s test sequence are shown for CVS PM in Figure 35 and for PFDS PM in Figure 36. Data are shown from all test labs, with concurrently sampled data from the CVS and PFDS shown.

From the CVS (Figure 35), profiles from the JRC, UTAC and AVL-MTC tests are relatively similar, with PM levels remaining relatively flat through the test sequence. The labs with the high tunnel backgrounds show different profiles. These results suggest that if any changes in emissions levels occur due to cycle-to-cycle variations or changes to DPF fill state occur with passive regeneration within individual cycles, CVS PM is insensitive to them. Alternatively, the results suggest that in this testing any effects that do occur have been masked by tunnel background levels even in the labs with low emissions.

Figure 35: CVS PM Emissions through the Test Sequence – All labs

From the partial flow system (Figure 36), PM levels were relatively consistent through the test sequence from all labs except UTAC where the tunnel background levels were higher. It is possible that the PM method detects a reduction in cycle PM between the cold and hot WHTC cycles, but as with the CVS data, the PFDS PM measurements in this testing were either insufficiently sensitive to detect changes in DPF fill state through the day’s test sequence, or the effects were masked by tunnel background levels.

Figure 36: PFDS PM Emissions through the Test Sequence – All labs

### Overview of PM Results

- CVS PM results, after elimination of the high tunnel background labs data, showed emissions levels of <6mg/kWh across all cycles

- Emissions levels from the PDTs were generally lower at 4mg/kWh or less from all cycles, with the exception of UTAC’s results which were in the range 4mg/kWh to 7mg/kWh. This may have been due to a higher PFDS tunnel background than other labs, but there is insufficient data to draw a firm conclusion

- CVS PM tunnel backgrounds were generally at the same level as samples except from ESC cycles, as ESC PM contains low volatility HCs which, once collected, remain on the filter. As a consequence, the CVS PM measurements in this testing only appeared capable of resolving engine emissions from the tunnel background for ESC tests.

- Tunnel background PM levels in PFDS systems were at the low end of levels seen from CVS systems. Even so in only one of 3 systems, in which tunnel background levels were determined, was it possible to discriminate data from any more cycles than the ESC.

- In the one PFDS system that enabled discrimination between tunnel background and sample levels, emissions from all cycles were <2mg/kWh and specifically ~1.2mg/kWh from the WHSC and <1mg/kWh from the weighted WHTC.

## Full Flow and Partial Flow PN

### PN – Tunnel background Levels

PN tunnel background levels appear to vary substantially between CVS systems in different laboratories, but appear highly similar between partial flow systems. While data are not available for tunnel backgrounds in all labs, comparisons of hot transient cycle results from different labs, where particle emissions are low, clearly shows the offsets due to tunnel backgrounds.

As Figure 37 shows, particle number emissions measured from the CVS from different labs across the hot WHTC can vary by a factor of 100 or more. In contrast (Figure 41), particle number emissions from PDTs sampled simultaneously to the CVS data appear to overlay.

JRC, AVL and Ricardo (RCE) also supplied particle number data from 30 minutes dilution system tunnel background acquisitions taken prior to the cold start WHTC on each test day. A typical tunnel background from each of the PFDS and CVS facilities at these labs was calculated to particles/kWh and the values obtained were compared. These data are shown in Figure 38.

The CVS tunnel background at Ricardo was >4x1010/kWh. This was ~60 times higher than the CVS tunnel background at AVL-MTC and ~140 times higher than the CVS tunnel background at JRC. These differences are in line with the differences in emissions levels seen from the hot start transient test shown in Figure 37, and indicate that from Ricardo and to a lesser extent AVL, the tunnel background predominates in the particle numbers measured on these tests.

The tunnel background levels from the 3 PFDS systems are almost identical: at <3x108/kWh. This indicates that the small variations in particle number levels seen in Figure 38 are probably related to real engine or DPF variability rather than tunnel background contributions.

As noted in Section 6.1.4, where PM tunnel background was discussed, the Ricardo facility is a working facility which had recently experienced testing on high bio-content fuels, work on active regeneration strategies for DPF regeneration and substantial amounts of non-DPF testing on low NOx calibration engines. An elevated PM tunnel background was seen from this lab and it is apparent the tunnel background also contributes substantially to PN.

Figure 37: Tunnel background Impacts PN Results – CVS Systems at 3 Labs

Figure 38: Tunnel background Effect on PN Results Minimal– PFDS Systems at 3 Labs

Figure 39: Particle Number Tunnel backgrounds – 3 Labs; CVS and PFDS Facilities

### PN - Repeatability

#### Transient Particle Production

Figure 40 to Figure 44 inclusive show real-time repeatability traces for cold WHTC, hot WHTC, WHSC and ESC tests. Each figure shows PN results from a lab with the highest CVS particle number tunnel background (upper), results from the lab with the lowest CVS PN tunnel background (middle) and typical results from a partial flow system (lower). The bottom and middle charts' data are drawn from the same test laboratory.

Data from the cold start WHTCs shown in Figure 40 (upper) covers approximately two orders of magnitude, but the range seen in the middle and lower figures covers more than 4 orders of magnitude. It is also clear that data in the top figure shows little similarity to the engine speed after the first 700s, but the middle and lower figures reflect changes in engine operation throughout the emissions cycle.

The high levels of tunnel background seen in the CVS (Figure 40, upper) do not have a substantial effect on repeatability, because the overall emissions levels from the cycle are dominated by those of the first 700s. Repeatability levels from the data shown, based upon mean cycle results, would be relatively similar to those seen for Figure 40, middle and lower.

The contrast between the upper and middle parts of Figure 40 is clear: while the profiles of emissions coincide at the peaks, the less transient parts of the emissions cycles can be masked if the tunnel background is high. Results from the lower figure are highly similar to those of the middle figure, indicating that there are not fundamental differences in either emissions levels or transient measurements from full and partial flow systems as long as tunnel background levels are similar.

The sensitivity of the measurement is obviously affected by the tunnel background, but as mentioned above, the dominance of the peaks seen in the first 700s makes even a high tunnel background almost irrelevant to results from the cold WHTC emissions cycle.

Results from the hot start WHTC (Figure 41) were substantially affected by the high CVS tunnel background. A comparison between high (upper figure) and low (middle figure) tunnel background CVS facilities shows that the transient traces from the high tunnel background lab bear little or no relationship to the transient events of the cycle and span less than a factor of 10. Conversely, the emissions of the low tunnel background CVS (middle figure) and from the partial flow system (lower figure) indicate transient events corresponding to engine operation and emissions ranges that span 2 to 3 orders of magnitude. Once again, results from the flow tunnel background CVS and partial flow systems appear very similar.

Peak emissions levels from the partial flow system (and low tunnel background CVS) were ~ 1000 times lower than the peaks from the cold start WHTC, but in the high tunnel background CVS, cold WHTC emissions only appeared to be 10 times higher than the hot start WHTC. Clearly, the hot WHTC repeatability from the high tunnel background CVS system reflects little more than the repeatability of the tunnel background levels.

Emissions from the WHSC cycle (Figure 42) were at a similar level to those seen from the hot WHTC (Figure 41), but the emissions profile from the high tunnel background CVS (upper figure) tracks the engine speed trace well. This suggests that the tunnel background is less significant during this cycle than in the hot WHTC, so running the previous two cycles in the test matrix may have had a ‘cleaning’ effect on the CVS.

From all three dilution tunnels, it appears that particle number emissions are very low from the WHSC until ~1200s. The exhaust temperature of the WHSC is sufficient to enable passive regeneration (where NO2 reacts with soot on the DPF to produce N2 and CO2) during ~60% of the emissions cycle (Figure 43, left).

It is possible that after ~1200s of the WHSC, the soot loading of the DPF has reduced sufficiently to have a substantial impact on filtration efficiency. The exhaust temperature has also increased to ~450°C at this point in the cycle (Figure 43, right) and this might lead to the thermal release of low volatility components from the exhaust system. Both these mechanisms could increase particle number emissions and variability.

The ETC cycle is tested following the WHSC in the test sequence. Results, emissions levels and repeatability from this cycle were very similar to those from the hot WHTC. The ETC runs at relatively low exhaust temperatures, so during this cycle additional soot is added to the DPF following passive regeneration in the WHSC. This may have the effect of increasing filtration efficiency for the start of the ESC.

Figure 40: Real-time Data – Cycle-to-cycle Repeatability, Cold WHTC

Figure 41: Real-time Data – Cycle-to-cycle Repeatability, Hot WHTC

Figure 42: Real-time Data – Cycle-to-cycle Repeatability, Hot WHSC

Figure 43: Passive Regeneration During the WHSC

Emissions during the 4 minutes of idle at the start of ESC cycle (Figure 44 upper, middle and lower) were very low and probably indistinguishable from tunnel background in all dilution systems. As with the WHSC cycle, most of the modal transitions of the ESC are visible even in the high tunnel background data (upper figure), but the range of emissions seen in this system (~100 from low to high) compares poorly with the 4 orders of magnitude from the low tunnel background CVS and PFDS and indicates lowered sensitivity of this measurement system.

The ESC shows the highest variability levels on a cycle-to-cycle basis, and these appear to worsen after 1000s of the cycle: at this point exhaust temperatures rise substantially, reaching >600°C after ~1300s. Passive regeneration in this cycle may eliminate the soot replaced during the ETC cycle and further eliminate soot from the DPF. Any variations in initial soot loading during the previous night's preconditioning may manifest as variability in the ESC.

Figure 44: Real-time Data – Cycle-to-cycle Repeatability, Hot ESC

#### Repeatability as CoV

Repeatability levels for all labs are given as single CoV values that express overall intra-lab variability for each emissions cycle (see Section 4).

Figure 45 shows the repeatability of the 5 test matrix cycles and the composite weighted WHTC result for the CVS-based and PFDS-based PN methods.

Three results are shown for each cycle, and these include:

- PMCVS1: All data from all labs (excepting tests excluded for technical reasons)

- PMCVS2: Outlier analysis iteration 1

- PMCVS3: Outlier analysis iteration 2

Outlier analyses excluded the CVS PN results from both Ricardo and EMPA on the first pass iteration excepting: cold WHTC data from both labs and ESC data from EMPA.

On the second pass analysis, the ETC data from the first batch of tests at JRC were excluded.

Considering all cycles (Figure 24), and following the outlier iterations, repeatability levels were broadly similar: CVS CoVs ranged from ~20% to ~60% and PFDS CoVs from ~20% to ~70%.

Focusing on the Euro VI legislative cycles in isolation shows that the CVS approach has better repeatability over the weighted WHTC (21.1% Vs. 22.8%) and over the WHSC (59.2% Vs. 74.4%) than the PFDS approach.

Figure 45: Repeatability of Particle Number Measurement Systems

### PN - Reproducibility

Reproducibility levels for all labs are given as single CoV values that express overall inter-lab variability from each emissions cycle (see Section 4).

Figure 46 shows the reproducibility of the 5 test matrix cycles and the composite weighted WHTC result for the CVS-based and PFDS-based PN methods.

The outlier analysis has a substantial effect on reducing the variability of the CVS system’s results (Figure 46 upper), but has little impact on the PFDS results, except in the ETC where the first results from JRC were eliminated from the dataset.

Considering just the Euro VI legislative cycles shows that the CVS approach has better reproducibility over the weighted WHTC (41.4% Vs. 45.8%) and over the WHSC (81.7% Vs. 86.3%) than the PFDS approach.

Figure 46: Reproducibility of Particle Number Measurement Systems

### PN – Emissions Levels

Particle number emissions from the various test cycles are shown for CVS measurements in Table 10 and Figure 47 and for PFDS measurements in Table 11 and Figure 49. These Tables and Figures include all test results including those rejected as outliers by the statistical analyses.

#### PN Emissions from CVS Systems

Particle number emissions from the cold WHTC cycle ranged by approximately an order of magnitude across all laboratories – from ~6 x1010/kWh to ~7x1011/kWh with the all-labs mean at ~4x1011/kWh.

Hot WHTC results ranged from 109/kWh (JRC 2nd campaign) up to ~5x1011/kWh at Ricardo, a difference from low to high of 500 times. The all labs mean, at ~6x1010/kWh, was substantially impacted by the labs with high emissions levels.

Table 10: Maximum and Minimum Ranges for CVS-measured PN

Weighted WHTC results, based upon the cold and hot WHTC data providing 10% and 90% contributions respectively, ranged from ~7x109/kWh to ~5x1011/kWh with an all labs mean of ~1011/kWh.

WHSC and ETC results reflected the results of the hot WHTC, with lowest emissions levels at JRC and highest at Ricardo. WHSC emissions levels ranged by more than 200 times: from 2x109/kWh up to ~4x1011/kWh and ETC ranged by a factor of >300, from ~109/kWh up to ~3..5x1011/kWh.

ESC results showed a narrower range than any of the emissions cycles except the Cold WHTC. Low to high range covered a factor of <30, from ~2x1010/kWh up to ~5.5x1011/kWh.

Figure 47: Maximum and Minimum Ranges for CVS-measured PN

As described in Section 4, statistical methods were used to eliminate individual PN data and a majority of data from several cycles at two specific laboratories in the dataset. Excluded data are:

- Ricardo CVS data from WHTC Hot, WHSC, ETC, ESC

- EMPA CVS data from WHTC Hot, WHSC, ETC

- ETC data from the first set of tests at JRC

The effects on the mean of means (the average PN emissions across all labs) of eliminating the outlier data is shown in Figure 48. The Cold WHTC data are unaffected since no data are excluded, but the hot WHTC result is substantially reduced, resulting in a weighted WHTC reduction of >50% to~5x1010/kWh. The WHSC result drops to <2x1010/kWh, the ETC to below 1010/kWh and the ESC to <7x1010/kWh.

Figure 48: CVS PN Results – Effect of Eliminating Outliers

#### PN Emissions from PFDS Systems

Particle number emissions from the cold WHTC cycle ranged by approximately an order of magnitude across all laboratories – from ~6x1010/kWh to ~7x1011/kWh with the all-labs mean at ~3.7x1011/kWh. From the PFDS, cold WHTC emissions levels were substantially higher than from any other cycles.

Table 11: Maximum and Minimum Ranges for PFDS-measured PN

Hot WHTC results ranged from <109/kWh (JRC 2nd campaign) up to ~1.6x1010/kWh, a difference from low to high of less than 20 times. The all labs mean was ~ 5x109/kWh, almost 2 orders of magnitude lower than the cold start emissions.

The low emissions levels measured from PFDS during the hot WHTC cycle resulted in a substantial reduction in the weighted WHTC result relative to the cold start cycle. Emission ranged from ~7x109/kWh to ~8x1010/kWh with the all labs mean at ~4.2x1010.

Emissions from the WHSC showed the greatest range, a factor of ~120, which reflects the substantial period of operation spent in passive regeneration during this cycle and the influence that has on DPF fill-state and filtration efficiency. Low-end emissions levels were close to 109/kWh increasing to 1.4x1011/kWh.

The range and absolute emissions levels from the ETC were similar to those seen from the hot WHTC: ~109/kWh up to ~2x1010/kWh with the all labs mean at ~6x109/kWh.

ESC data from Ricardo were highest of all the labs, but not identified as an outlier. Emissions ranged from ~2x1010 to ~3x1011/kWh with the all labs mean at just below 8x1010/kWh.

Figure 50 shows the mean of all labs PN emissions from PFDS systems. Data are shown for all emissions cycles.

Figure 49: Maximum and Minimum Ranges for PFDS-measured PN

Figure 50: All Labs’ Mean of Means PN Results - PFDS Systems

#### PN Emissions Levels Overview

In both CVS and PFDS cases, and considering mean of means data, emissions were highest from the cold start WHTC at ~4x1011/kWh. At this level of emissions, contributions from the tunnel background, even from labs with very high tunnel backgrounds, do not have a substantial impact on emissions.

Lowest emissions were observed from hot cycles which do not have substantial periods of passive regeneration: hot WHTC and ETC. These cycles showed emissions levels of 5 - 6x109/kWh from the PFDS and 8 - 9 x109/kWh from the CVS once outlier laboratories were excluded. Laboratories considered to be outliers reported emissions levels from these cycles substantially above 1011/kWh.

Weighted WHTC results were of the order 4 - 5x1010/kWh from the PFDS and CVS (outliers excluded) and ~1011/kWh from the CVS when all labs data were considered.

ESC and WHSC cycles results were generally more variable than the hot start ETC and WHTC due to the presence of passive regeneration during these cycles. Passive regeneration may reduce filtration efficiency by reducing or removing the filter cake but high temperatures may also liberate low volatility HCs which contribute to solid particles: both of these result in higher PN emissions from the ESC and WHSC than from the ETC and hot WHTC. The ESC cycle, which has a 2 minutes period of operation at full load, and other modes with very high exhaust temperatures, sees a higher contribution of low volatility HC ‘solid particles’ than the lower temperature WHSC. Emissions from the WHSC were around 2-3x1010/kWh from PFDS and CVS (outliers excluded) and 6 - 8x1010/kWh from the ESC with outliers excluded from the CVS data.

#### Euro VI Certification

Certification testing for Euro VI will include both WHTC and WHSC. From the emissions levels seen in this study a limit set mandating PFDS and excluding CVS could be lower than a limit set that permits either dilution system. If either dilution system is permitted, the test laboratory or OEM would face substantially higher risk of non-compliance when using a full flow system, unless that CVS was well characterised beforehand.

#### Achievable PN Limit for this Engine

Across this test programme the highest single PN emissions result seen from any cycle across all labs was 7.4x1011/kWh from a cold WHTC. If this engine and its DPF are considered to have representative PN emissions and both CVS and PFDS dilution approaches are used, a PN limit of 8x1011/kWh would be achievable for all emissions cycles tested.

## Gaseous Emissions

### General Observations

Gaseous regulated emissions results were generally more repeatable when measured directly from raw exhaust than when measured as dilute emissions. Substantial differences in raw vs. dilute emissions levels also exist, even for ‘high emissions’ gases such as CO2. Evaluations of these differences were not the focus of this programme, but data from this work will be made available for consideration in relevant studies.

Some labs had high levels of CO and HC backgrounds even when these emissions were measured from raw exhaust. This presents a challenge for the measurement of, in particular, carbon monoxide and hydrocarbons at low levels.

It is clear that few labs regularly measure ‘bagged emissions’ from HD testing. This approach of accumulating a dilute sample permits a true, simultaneously sampled, background to be quantified and subtracted. Even so, HC and CO emissions are so low from modern engines that even this approach, since it samples from diluted exhaust, may lack accuracy.

### Inter-Laboratory Variability

During the ILCE_LD, the inter-laboratory variations of regulated gases over the cold start NEDC were found to be ≤5% for CO2, 35-50% for HC and 10-30% for NOx and >40% for CO.

Considering raw emissions measurements during the ILCE_HD, inter-lab variations were at broadly similar levels from the Cursor 8 engine, with some emissions cycles showing lower variations than others. Highest variations were seen from CO and HC, where actual g/kWh emissions levels were low. Ranges of variation encompassing all test cycles are given below:

- CO2 emissions across the labs: inter-lab variation range (all test cycles) 6% - 13%

- CO emissions across the labs: inter-lab variation range (all test cycles) 63% -91%. CO emissions range from 13mg/kWh to 400mg/kWh

# Comments

Expand and explain which cycles gave best and worst repeatability

# Endnotes

References

Conclusions on Improving Particulate Mass Measurement Procedures and New Particle Number Measurement Procedures Relative to The Requirements of The 05 Series of Amendments to Regulation No. 83. http://www.unece.org/trans/doc/2004/wp29grpe/TRANS-WP29-GRPE-48-inf11r1e.pdf

A working document for the UK Department for Transport presenting an updated and restructured version of Regulation No. 49, Annex 4. http://www.unece.org/trans/doc/2004/wp29grpe/PMP-2004-13-03e.doc

Regulation 83: Vehicle Emissions Legislation; Light-duty vehicles and PCVs (GVW < 3.5 tonnes). http://www.unece.org/trans/main/wp29/wp29regs/r083r3e.pdf

Regulation 43: Vehicle Emissions Legislation; Heavy Duty Vehicles (GVW >3.5 tonnes). http://www.unece.org/trans/main/wp29/wp29regs/r049r3e.pdf

UN-GRPE PMP Phase 3 Inter-laboratory Correlation Exercise: Framework and Laboratory Guide, A Document For The UK Department for Transport. RD 04/80801.5 Jon Andersson and David Clarke

http://ies.jrc.ec.europa.eu/uploads/fileadmin/Documentation/Reports/Emissions_and_Health/EUR_2006-2007/EUR_22775_EN.pdf

Giechaskiel B., Dilara P., Andersson J. (2008). Particle Measurement Programme (PMP) light-duty inter-laboratory exercise: Repeatability and reproducibility of the particle number method. Aerosol Sci. Technology, 42, 528-543

Giechaskiel B., Dilara P., Sandbach E., Andersson J. (2008). Particle Measurement Programme (PMP) light-duty inter-laboratory exercise: Comparison of different particle number measurement systems. Measurement Sci. Technology, 19, 095401

http://www.unece.org/trans/main/wp29/wp29regs/r083r3a2e.doc

Exploratory experiments SAE 2009-01-1767

Particle Measurement Programme (PMP) Heavy-Duty (HD) Inter-laboratory Exercise. Exploratory work at JRC (Oct’ 07 -Dec’07). B. Giechaskiel, S. Alessandrini, F. Forni, P. Sinuez, M. Carriero, G. Martini

UN-GRPE PMP Phase 3 Inter-laboratory Correlation Exercise: Updated Framework and Laboratory Guide for Heavy Duty (HD) Engine Testing A Document For The UK Department for Transport. Ricardo Document  RD 07/578701.2a. Jon Andersson and David Clarke

ISO 16183:2002 Heavy duty engines -- Measurement of gaseous emissions from raw exhaust gas and of particulate emissions using partial flow dilution systems under transient test conditions. http://www.iso.org/iso/catalogue_detail.htm?csnumber=32152

http://www.horiba.com/automotive-test-systems/products/emission-measurement-systems/dilution-sampling-systems/details/mdlt-1300t-876/

AVL SPC 472 Smart Sampler: http://www.avl.com/wo/webobsession.servlet.go/encoded/

YXBwPWJjbXMmcGFnZT12aWV3JiZub2RlaWQ9NDAwMDIwMTc0.html

http://www.controlsistem.it/data/automotive/testrig/pss20.swf

http://sierrainstruments.com/emissions/bg3.html

UK Particulate Measurement Programme (PMP): A Near US 2007 Approach To Heavy Duty Diesel Particulate Measurements - Comparison With The Standard European Method. Jon D Andersson; David P Clarke; James A Watson. SAE 2004-01-1990

Wei Q., Oestergaard K., Porter S., Asano I, Adachi M. and Rahman M., 2006. Real-time Measuring System for Engine Exhaust Solid Particle Number Emission – Design and Performance. SAE 2006-01-0864

Johnson K. C., Durbin T. D., Jung H., Chaudhary A., Cocker D. R., Herner J. D., Robertson W. H., Huai T., Ayala A. and Kittelson D., 2009. Evaluations of the European PMP methodologies during an-road and chassis dynamometer testing for DPF equipped heavy-duty diesel engines. Aerosol Science and Technology, 43, 962-969.

Johnson, Kent C., Thomas D. Durbin, Heejung Jung, Ajay Chaudhary, David R. Cocker III, Jorn D. Herner, William H. Robertson, Tao Huai, Alberto Ayala, and David Kittelson, 2009. Evaluation of the European PMP Methodologies during On-Road and Chassis Dynamometer Testing for DPF Equipped Heavy Duty Diesel Vehicles, Aerosol Science and Technology, 43:962–969, 2009.

Johnson, Kent C., Thomas D. Durbin, Heejung Jung, Ajay Chaudhary, David R. Cocker III, Jorn D. Herner, William H. Robertson, Tao Huai, Alberto Ayala, and David Kittelson, 2009. Evaluation of the European PMP Methodologies during On-Road and Chassis Dynamometer Testing for DPF Equipped Heavy Duty Diesel Vehicles, Aerosol Science and Technology, 43:962–969, 2009.

Hines W., Montgomery D., Goldsman, D. and Borror C., Probability and Statistics in Engineering. 4th edition, John Wiley and Sons, 2003.

Armstrong S., 2007. Significance tests harm progress in forecasting. International Journal of Forecasting, 23 321-327.

Mamakos A., Ntziachristos L. and Samaras Z., 2004. Comparability of particle emission measurements between vehicle testing laboratories: a long way to go. Measurement Science and Technology, 15, 1855-1866.

Penetration Calibration and Verification for the Solid Particle Counting System with Polydisperse and Monodisperse Particles. Qiang Wei, Mike Akard, Ichiro Asano , Montajir Rahman. SAE 2008-01-1178

Real-time measuring system for engine exhaust solid particle number emission – Performance and Vehicle tests.Qiang Wei, Asano Ichiro, Masayuki Adachi, Rahman M. Montajir, Takeshi Kusaka, Yuichi Goto. SAE 2006-01-0865

Calibration of Condensation Particle Counters. Wei Liu, Brian L. Osmondson, Oliver F. Bischof, and Gilmore J. Sem. SAE 2005-01-0189

Personal communication between J.Andersson and Hans-Georg Horn (TSI Europe)

Final Report For DfT On Testing In The PMP Phase 3 Heavy-Duty Validation Exercise. Ricardo Document 09_110201_2

Particle Measurement Programme (PMP) Light-Duty Inter-laboratory Exercise: Repeatability and Reproducibility of the Particle Number Method. Barouch Giechaskiel, Panagiota Dilara, Jon Andersson. Aerosol Science and Technology, Volume 42, Issue 7 July 2008 , pages 528 - 543

# Text boxes

e

d

c

b

a

e

d

c

b

a

a

Comparison of Tunnel Background and Sample Filter Masses at EMPA

Comparison of Tunnel Background and Sample Filter Masses at JRC

Comparison of Tunnel Background and Sample Filter Masses at EMPA

Comparison of Tunnel Background and Sample Filter Masses JRC

Comparison of Tunnel Background and Sample Filter Masses at EMPA

SPCS20 – Low PFDS Tunnel Background

SPCS19 – Low CVS Tunnel Background

SPCS19 – High CVS Tunnel Background

SPCS20 – Low PFDS Tunnel Background

SPCS19 – Low CVS Tunnel Background

SPCS19 – High CVS Tunnel Background

SPCS20 – Low PFDS Tunnel Background

SPCS19 – Low CVS Tunnel Background

SPCS19 – High CVS Tunnel Background

SPCS20 – Low PFDS Tunnel Background

SPCS19 – Low CVS Tunnel Background

SPCS19 – High CVS Tunnel Background